PROPERTIES AND OPTIMIZATION OF RESPIRATORY NAVIGATOR GATING FOR SPIRAL CINE DENSE CARDIAC MRI

DISSERTATION

A dissertation submitted in partial fulfillment of the requirements of the degree of Doctor of Philosophy in the College of Engineering

at the University of Kentucky

By

Sean Michael Hamlet

Lexington, Kentucky

Co-Directors: Dr. Brandon Fornwalt, Professor of Electrical Engineering and Dr. Kevin Donohue, Professor of Electrical Engineering

Lexington, Kentucky

Copyright Oc

2017

Sean Michael Hamlet 2017

ABSTRACT OF DISSERTATION

PROPERTIES AND OPTIMIZATION OF RESPIRATORY NAVIGATOR GATING FOR SPIRAL CINE DENSE CARDIAC MRI

Cardiac magnetic resonance  (CMR)  can non-invasively assess  heart function. Displacement  encoding with  stimulated echoes  (DENSE)  is an advanced CMR imaging technique that measures tissue displacement  and can be used to quantify cardiac mechanics  (e.g.   strain and torsion).   When combined with  clinical risk factors, cardiac mechanics  have  been shown to be  better predictors of mortality than traditional  measures  of function.   End-expiratory breath-holds are typically used to minimize respiratory motion artifacts.  Unfortunately, requiring subjects to breath-hold introduces limitations  with  the  duration  of  image acquisition and quality of data acquired, especially in patients with limited breath-holding ability. Thus, DENSE acquisitions often require respiratory navigator gating, which works by measuring the diaphragm during normal breathing and only acquiring data when the diaphragm is within a pre-defined acceptance window.

Unfortunately,  navigator gating results in  long scan durations due to inconsistent  breathing patterns.  Moreover, the respiratory navigator configuration directly affects image quality. Scan duration and image quality need to be optimized in order to improve the clinical utility  of DENSE. Thus, the goal of this project was to optimize those parameters.  To accomplish this goal, we set out to 1) understand how respiratory gating affects measures  of cardiac mechanics,  2) determine the optimal respiratory navigator configuration, and 3) reduce  scan duration by using an interactive videogame.

Aim 1 of this project demonstrated that the variability in torsion, but not strain, could be significantly reduced through the use of a respiratory navigator compared to traditional  breath-holds. Aim 2 of this project  demonstrated that,  among the configuration options, the dual-navigator configuration resulted in the best image quality compared to gold standard traditional breath-holds, but also resulted in the longest scan duration.   Aim  3 demonstrated that  using an interactive  breathing- controlled videogame during CMR can significantly reduce scan duration compared to traditional free-breathing.

In summary, respiratory navigator gating with DENSE 1) reduces the variability in measured LV torsion, 2) results in the best image quality with the dual-navigator configuration, 3) results in significantly shorter scan durations through the use of an interactive videogame. Selecting the optimal navigator configuration and using an interactive videogame can improve the clinical utility  of DENSE.

KEYWORDS:  Respiratory Navigator Gating,  Cardiac Magnetic Resonance Imaging, Displacement Encoding with  Stimulated Echoes,  Cardiac Mechanics, Interactive Videogame

Sean Michael Hamlet

Today’s Date

Date

PROPERTIES AND OPTIMIZATION OF RESPIRATORY NAVIGATOR GATING FOR SPIRAL CINE DENSE CARDIAC MRI

By

Sean Michael Hamlet

Dr. Brandon Fornwalt, MD, PhD Co-Director of Dissertation

Dr. Kevin Donohue, PhD Co-Director of Dissertation

Dr. Caicheng Lu, PhD Director of Graduate Studies

(Date) Date

To Regina

Once again, I could not have done this without your love and support.

Thank you for your continuing positivity and encouragement.

ACKNOWLEDGMENTS

Although this research was completed with my individual effort, its progress could not have been completed without the efforts of several people.

Paragraph 2

TABLE OF CONTENTS Acknowledgments . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	iii List of Tables . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . viii List of Figures . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	 x

Chapter 1:	Background . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	1

1.1	Heart Disease . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	1

1.2	Standard  Cardiac Magnetic  Resonance Imaging  (MRI) and Traditional

Measures of Cardiac Function . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	1

1.3	Advanced Measures of Function: Cardiac Mechanics . . . . . . . . . . . . . . . . . . . . .	2

1.4	Displacement Encoded Cardiac MRI . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	4

1.5	Respiratory Motion and Blurring . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	4

1.6	Respiratory Navigator Gating . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	4

1.7	Dissertation Outline . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	5

Chapter	2:	Effects	of	Patient-Specific	Variability 	in 	Inconsistent End-Expiratory Diaphragm Position On the Quantification of Left Ventricular Cardiac Strains . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	7

2.1	Synopsis . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	7

2.2	Background . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .	9

2.3	Methods . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  10

2.3.1	Subjects . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  10

2.3.2	Quantification of Inconsistent End-Expiratory Positions . . . . . . . . .  11

2.3.3	DENSE Acquisition . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  11

2.3.4	DENSE Post-Processing . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  13

2.3.5	Statistics . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  14

2.4	Results . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  15

2.4.1	Inconsistent End-Expiratory Positions . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  16

2.4.2	Differences and Variability in Peak Strains . . . . . . . . . . . . . . . . . . . . . . . .  16

2.5	Discussion . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  17

Chapter	3:	Using	a	Respiratory	Navigator	Reduces	Variability 	when

Quantifying Left Ventricular Torsion . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  22

3.1	Synopsis . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  22

3.2	Background . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  25

3.3	Methods . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  26

3.3.1	LV Motion Quantification . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  27

3.3.2	Experiment 1: Enforced End-Expiratory Variability . . . . . . . . . . . . . .  28

3.3.3	Experiment 2: Natural End-Expiratory Variability . . . . . . . . . . . . . . .  29

3.3.4	Statistics . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  31

3.3.5	Theoretical Sample Size Calculation . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  32

3.4	Results . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  32

3.4.1	Inconsistent End-Expiratory Positions . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  33

3.4.2	Torsion . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  33

3.4.3	Theoretical Sample Sizes . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  35

3.5	Discussion . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  37

3.5.1	Limitations and Future Directions . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  41

3.6	Conclusion . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  42

Chapter 4:	Optimal Respiratory Navigator Configuration . . . . . . . . . . . . . . . . . . . . . . . .  43

4.1	Synopsis . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  43

4.2	Introduction . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  45

4.3	Methods . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  46

4.3.1	Subjects . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  46

4.3.2	Image Acquisition . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  46

4.3.3	Navigator Feedback . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  49

4.3.4	DENSE Post-Processing . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  50

4.3.5	Analysis . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  50

4.3.6	Comparison of Acquisition Configurations . . . . . . . . . . . . . . . . . . . . . . . . .  51

4.4	Results . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  52

4.4.1	Average Peak Strains . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  52

4.4.2	Signal-to-Noise Ratio . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  55

4.4.3	Navigator Efficiency . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  55

4.5	Discussion . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  56

Chapter  5:	Interactive   Feedback Game Designed to  Improve   Navigator

Efficiency . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  63

5.1	Synopsis . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  63

5.2	Background . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  65

5.3	Methods . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  66

5.3.1	Feedback videogame . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  66

5.3.2	Participants . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  68

5.3.3	Imaging . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  68

5.3.4	Calculation of cardiac strains from DENSE . . . . . . . . . . . . . . . . . . . . . . .  70

5.3.5	Analysis . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  71

5.3.6	Training . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  71

5.3.7	Statistics . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  72

5.4	Results . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  73

5.4.1	Navigator Efficiency . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  73

5.4.2	SNR . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  73

5.4.3	Heart rate . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  76

5.4.4	Strain . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  76

5.5	Discussion . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  78

5.5.1	Navigator Efficiency . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  78

5.5.2	SNR . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  79

5.5.3	Heart Rate . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  80

5.5.4	Strains . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  81

5.5.5	Clinical Implications . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  81

5.5.6	Comparison with Previous Work . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  82

5.5.7	Limitations . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  82

5.5.8	Conclusion . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  85

Appendices . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  86

Chapter A: 	Appendix Differences  in  Segmental Strain  Between Different

Acceptance Window Positions . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  87

Chapter B: 	Appendix Segmental Strains for Navigator Gating . . . . . . . . . . . . . . . . . .  90

B.1  Peak Segmental Strains for Navigator Gating and Breath-holds . . . . . . . . .  90

B.2  Segmental Strain Agreement Between Navigator Gating and Breath-

holds . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  90

Chapter C:	Appendix Off-Scanner Training Protocol for Feedback Game . . . . . . .  92

C.1  Training Protocol . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  92

C.2  Survey  Responses . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  92

LIST OF TABLES

2.1	Global and regional peak strains (mean	standard deviation) from the three acceptance window positions (minimum, middle, and maximum) for

all subjects combined ............................................................................  17

2.2	Standard	deviation	of	global	and	regional	peak	strains	across inconsistent  (maximum, middle, minimum) and consistent  (middle and repeated middle) acceptance window positions.........................................  18

2.3	Power analyses  for studys ability  to detect a difference  in global and regional strain between different end-expiratory positions ......................... 18

3.1	Mean (± standard deviation) of torsion across the volunteers within each experiment ...........................................................................................  36

3.2	Mean (± standard deviation) of torsion across the volunteers within each

experiment ...........................................................................................  38

4.1	Average strains for different acquisition techniques...................................  53

4.2	CI  Results for  Differences in  Strain  Between Navigator  Gating  and

Breathhold DENSE ...............................................................................  53

4.3	Signal-to-noise ratios for different navigator gating configurations in adults

and children .........................................................................................  55

4.4	Navigator efficiencies for different navigator gating configurations in adults

and children .........................................................................................  56

5.1	Average Heart Rate for Off-scanner Trained and Untrained groups ............  77

5.2	Global peak strain results for NF and FG scans ....................................... 78

A.1  Segmental circumferential strain (%, mean ± standard deviation) from the three acceptance window positions (minimum, middle, and maximum) for

all subjects combined ............................................................................  88

A.2  Segmental radial strain (%, mean ± standard deviation) from the three

acceptance window positions (minimum, middle, and maximum) for all subjects combined .................................................................................  89

B.1  Segmental strain results for navigator gating and breath-holds in adults ....  91

B.2  Segmental strain agreement  between navigator gating and breath-holds

from spiral cine DENSE ........................................................................  91

LIST OF FIGURES

1.1	Measuring cardiac strains dramatically improves  the ability  to predict mortality ..............................................................................................	2

1.2	Cardiac strain is a measure of how small segments of the  myocardium shorten or lengthen during contraction and relaxation ..............................	3

2.1	During respiration, diaphragm motion causes  the heart to translate a significant distance while the imaging plane remains fixed......................... 10

2.2	Respiratory navigator gating ..................................................................  11

2.3	A respiratory navigator was used to measure end-expiratory positions to define	the	patient-specific	minimum,	middle,	and	maximum

end-expiratory positions ........................................................................   12

3.1	Computation of LV torsion from basal and apical images .......................... 26

3.2	Real time images of the diaphragm as it translates during a respiratory

cycle ....................................................................................................  27

3.3	Respiratory navigator gating ..................................................................  29

3.4	Measured end expiratory diaphragm  positions were used to define subject specific maximum, middle, and minimum end expiratory positions ............  30

3.5	The nine possible torsion permutations  were constructed from three basal

and three apical images .........................................................................  30

3.6	Inconsistent end expiratory positions  across ten consecutive breath holds

in patients and healthy controls..............................................................   34

3.7	DENSE images from a representative  subject show the relative  twist differences between the basal and apical images at end systole .................. 35

3.8	Variability of torsion due to enforced inconsistent end-expiratory positions

versus the subject specific range of end-expiratory position ....................... 36

3.9	Variability of torsion due to naturally inconsistent end-expiratory positions

versus the standard deviation of end-expiratory position...........................  37

4.1	Different navigator gating configurations  used to acquire DENSE image data.....................................................................................................  47

4.2	Theoretical example to demonstrate the disadvantages of single- and dual- navigator gating configurations...............................................................   48

4.3	Bland-Altman plots of average peak circumferential (Ecc), radial (Err), and  longitudinal (Ell)  strains for retrospective, prospective,  and dual navigator gating vs breath-hold .............................................................. 54

5.1	Feedback videogame ..............................................................................  67

5.2	MRI Feedback Setup .............................................................................  68

5.3	Navigator Efficiency for Feedback Game vs No Feedback .......................... 74

5.4	Average navigator efficiency for Off-scanner Trained and Untrained groups	75

5.5	SNR for all No Feedback and Feedback Game images ............................... 75

5.6	Mean and standard deviation of heart rate for No Feedback and Feedback

Game acquisitions .................................................................................  76

C.1  Survey Question 1 .................................................................................  93

C.2  Survey Question 2 .................................................................................  94

C.3  Survey Question 3 .................................................................................  95

C.4  Survey Question 4 .................................................................................  96

C.5  Survey Question 6 .................................................................................  98

C.6  Survey Question 7 .................................................................................  99

CHAPTER 1

BACKGROUND

1.1  Heart Disease

Heart disease is the leading cause of death for both adult men and women [1]. The term ”heart  disease” can refer to conditions that  involve  narrow or blocked blood vessels, which can lead to a heart attack, or conditions that affect the heart muscle, valves, or rhythm, which can lead to inefficient pumping and heart failure [1]. In addition to adults, there are children who are born with congenital heart disease (CHD), heart defects that can be present at birth, and it’s a growing problem that affects over 2 million people in the US [2]. As surgical and treatment techniques have improved, children with CHD are living to adulthood. For both adults and children, in order to develop improved techniques for treatment and therapy, heart disease and cardiac function need to be accurately  monitored.

1.2  Standard 	Cardiac 	Magnetic 	Resonance	Imaging 	(MRI)	and

Traditional Measures of Cardiac Function

Magnetic  resonance imaging (MRI) is a non-invasive non-ionizing medical imaging technique that is used to assess the function and health of internal structures of the human body. MRI involves the use of strong magnetic fields and radio waves along with the natural hydrogen nuclei, which are abundant in human tissue, to generate images. Thus, MRI can be used to non-invasively assess the cardiac function and health.

Non-invasive imaging has become standard protocol for diagnosis, prognosis, and management of heart diseases  and for monitoring cardiac health.   Traditional measures  of cardiac function, such as  ventricular volumes,  ventricular mass, and ejection fraction, can be derived from standard cardiac MRI. Whole heart function

χ2 (for predicting mortality)

40 	p<0.001

30 	p=0.040

Clinical

Clinical + Ejection Fraction

20 	Clinical + Longitudinal Strain

10

0  *Data from Stanton et al; Circ Imaging 2009; 2:356-64

Figure 1.1: Measuring cardiac strains dramatically  improves the ability to predict mortality.

is typically assessed with these traditional metrics, but unfortunately, they may not contain enough information to explain the complex nature of some heart diseases. Moreover,  there is a growing body of evidence that suggests that, when combined with  clinical  risk  factors (e.g.    hypertension), advanced measures  of  cardiac mechanics  (e.g.   cardiac strains and torsion)  are better predictors of mortality compared to traditional measures [3] (Figure 1.1).

1.3  Advanced Measures of Function:  Cardiac Mechanics

Cardiac mechanics, such as strain and torsion,  measure the deformation of the heart as it contracts and relaxes throughout the cardiac cycle. Commonly  measures directions include circumferentially, radially, and longitudinally (Figure 1.2). Torsion measures the twisting motion of the heart along the longitudinal axis of the heart throughout the cardiac cycle. Cardiac mechanics can be quantified from analyzing the motion of small regions of the heart, which can be completed by using an advanced imaging technique called spiral cine Displacement ENcoding with Stimulated Echoes (DENSE).

A: Cardiac Strain

Circumferential Strain (%)

−

−

Figure  1.2: 	Cardiac  strain  is a measure of how small segments of the myocardium shorten or lengthen during contraction and relaxation.

1.4  Displacement Encoded Cardiac MRI

Spiral cine Displacement  ENcoding with  Stimulated Echoes  (DENSE) is an advanced cardiac magnetic resonance imaging technique that  directly encodes  the displacement of the myocardial tissue into the phase of the MR signal [4]. Because of its quantitative  nature, as opposed  to qualitative  such as  in standard cardiac MRI,  it allows for simple  and accurate quantification of cardiac mechanics.  In addition, DENSE has good spatial resolution and good reproducibility [5, 6].

1.5  Respiratory Motion  and Blurring

As with most cardiac MRI techniques,  images are typically acquired using end- expiratory breath-holds. Breath-holds are used to suspend respiration so the bulk motion of the heart is minimized during imaging. Respiratory compensation is in order to reduce motion blurring in images (Figure XX).  DENSE acquisitions are generally performed using end-expiratory breathholds ( 1520 s in duration) [cite 410]; however, this approach is constrained by the patients ability to breath- hold, which is limited in young subjects and many stages of advanced heart disease.  Furthermore, short acquisitions preclude the ability  to capture more robust data, such as three- dimensional (3D)DENSE,7,11,12 or high resolution imaging.13. In order to overcome this time limitation, a respiratory navigator has been used which allows the subject to breathe freely throughout image acquisition [cite].

1.6  Respiratory Navigator  Gating

Respiratory navigator gating works by measuring the diaphragm position during normal breathing and only acquiring data when the diaphragm is within a pre-defined acceptance window (Figure XX).  The trade-off of navigator gating is significantly increased scan duration because of poor navigator efficiency. For example, previous CMR studies have reported respiratory navigator efficiencies of 20 to 45% in adults

[47]. This poor navigator efficiency lengthens the duration of currently used clinical imaging and limits clinical feasibility of emerging advanced imaging techniques.

Now talk about how this time limitation can be fixed

Navigator efficiency is typically poor because breathing patterns can be erratic [810] and the patient  is gener- ally  unaware of the desired acceptance  window location.  Providing the patient  with  visual feedback of the dia- phragm position during CMR (navigator feedback)has been shown to improve breathing consistency and scan efficiency  in adults [5, 8].   For example, studies have  shown efficiency improvements up to 29

Moreover,  there are different  navigator configurations  possible, which all affect navigator  efficiency and  image quality   due  to  their  distinct  advantages and disadvantages.

Previous studies using navigator-gated DENSE have reported using a prospective single navigator configura- tion.7,12 However, there has been no formal comparison of the available navigator configurations. Moreover, the accuracy of derived cardiac mechanics and overall image quality for these navigator configurations compared with breathhold acquisitions  as a reference standard are largely unknown. The purpose of this study was to determine the optimal configura- tion of respiratory navigator gating for the quantification of left ventricular strain using spiral cine DENSE MRI

1.7  Dissertation Outline

The goal of this project was  to optimize respiratory navigator gating, which would improve the clinical utility  of DENSE. To accomplish this goal, we set out to

1) understand how respiratory gating affects measures  of cardiac mechanics,  2) determine the optimal respiratory navigator configuration, and 3) improve navigator efficiency, which reduces scan duration, by using an interactive breathing-controlled videogame during cardiac MRI.

In  Chapters 2 and 3,  we  address the  effects  of inconsistent  end-expiratory

diaphragm position  between breath-holds  and  respiratory  navigator  gating  on DENSE-derived cardiac mechanics, such as left ventricular strain and torsion.  In Chapter  2,  we  learn that  cardiac strain  is  insensitive  to  normal  changes in end-expiratory position between breath-hold DENSE acquisitions. In Chapter 3, we discover that  use of a respiratory navigator has the ability  to significantly reduce the variability  of cardiac torsion and thus the sample size needed to detect  small changes in torsion. The conclusions of the studies performed for Chapters 2 and 3 demonstrate the importance of employing a respiratory navigator or some form of consistent respiratory compensation for future studies.

In Chapter 4, we address the optimal navigator configuration.

CHAPTER 2

EFFECTS OF PATIENT-SPECIFIC VARIABILITY IN INCONSISTENT END-EXPIRATORY DIAPHRAGM POSITION ON THE QUANTIFICATION OF LEFT VENTRICULAR CARDIAC STRAINS

Adapted from...

2.1  Synopsis

Purpose: To determine if normal inconsistency in end-expiratory diaphragm position between separate image acquisitions significantly affects estimates of cardiac strains. Materials  and Methods:  We enrolled 17 subjects, including seven patients with heart disease.  For each subject, we measured the range of end-expiratory positions during 10 separate breath-holds. The imaging protocol comprised two  ventricular long-axis and three short-axis slices  of navigator-gated 2D cine displacement encoding  with  stimulated echoes  (DENSE) cardiac magnetic resonance  (MR).  To simulate end-expiratory position inconsistency, DENSE images were each acquired at the patient-specific minimum, middle, and maximum end-expiratory positions; a repeated acquisition at  the  middle  position  was   used to  quantify  variability independent   of  end-expiratory differences.      Differences and  variability   of  left ventricular peak strains were compared using analysis of variance and Students t-test.

Results:    The  range of  end-expiratory positions across 10  breath-holds was

10 ± 4 mm. There were no significant differences in global or regional peak radial, circumferential, or longitudinal strains measured at the different end-expiratory positions (p = 0.17–0.98). In general, there were also no differences in variability in global  or  regional peak strains  between inconsistent   (minimum,  middle,  and

maximum) and consistent  (two acquisitions from middle position) end-expiratory positions (p = 0.10–0.95). With  at least 80% power, the study had an ability  to detect global differences of 4.7%, 1.0%, and 1.7% (absolute) between end-expiratory positions for radial, circumferential, and longitudinal strains, respectively. Conclusion:  Measurements  of left ventricular peak strains with  DENSE cardiac MR are relatively insensitive to normal changes in end-expiratory position between separate image acquisitions.

Keywords: Cardiac Strains, Breath-holds, DENSE, Respiratory Navigator Gating

2.2  Background

Cardiac strains describe the deformation of myocardial tissue during contraction and relaxation.    Measures  of  cardiac strains have  been shown to  be  superior predictors of outcomes, such as  mortality,  compared to  traditional  measures  of cardiac function  or  traditional   clinical  risk  factors  alone [3].     Imaging  can non-invasively assess  cardiac strains using echocardiographic  techniques such as speckle tracking [7] and cardiovascular  magnetic resonance (MR) techniques such as myocardial feature tracking [8], myocardial tissue tagging [9, 10], phase  velocity mapping [11], strain  encoding [12], and displacement  encoding with  stimulated echoes (DENSE) [4, 13].

Peak    strains     vary     longitudinally     throughout     the     left     ventricle [14, 15, 16, 17, 18, 19, 20].  For example, previous studies have shown that  left ventricular radial, circumferential, and longitudinal strains vary between  the base and    apex   by    up    to    14%,    5%,    and    5%   (absolute),    respectively [14, 15, 16, 17, 18, 19, 20].    Cardiac MR  images are often  acquired during end-expiratory breath-holds to minimize respiratory motion artifacts. However, it is often difficult to achieve consistency in end-expiratory diaphragm position between successive breath holds, and variations of 4 to 13 mm are normal [21, 22, 23, 24, 25]. Inconsistent  end-expiratory positions  will  impact the position of the heart with respect to  the imaging plane (Figure 2.1).   For example, previous studies have reported short-axis and long-axis through-plane  displacements of up to 14 mm due to displacement  of diaphragm position between breath-holds [26, 27], and other studies have reported that the superior/inferior position  of the heart can displace

55-92% of the displacement of the diaphragm position [28, 29]. Because peak strains vary throughout the left ventricle, we hypothesize that translation of the heart with respect to the imaging plane to result in differences and variability  in measured strains.

End-Expiration	Mid-Inspiration	End-Inspiration

Cardiac Imaging Plane Original Diaphragm Position Current Diaphragm Position

Figure  2.1: 	During  respiration,  diaphragm  motion  causes the  heart  to translate a significant distance while the imaging plane remains fixed.

To our knowledge, no study has evaluated  the  sensitivity of cardiac strains to natural end-expiratory position variability.   This is an important  knowledge  gap, especially since the use of cardiac strains is increasing dramatically both in research and clinical  practice.   The purpose  of this  study  was  to  determine if  normal inconsistency  in end-expiratory position significantly affects the quantification of cardiac strains and therefore results in higher variability in measured cardiac strains compared to strains measured at a consistent end-expiratory position.

2.3  Methods

2.3.1 Subjects

The study protocol was approved  by the local Institutional  Review Board. Ten healthy volunteers with no known cardiovascular  disease or chronic illnesses and 7 patients  with   a  history  of  heart  disease (known  diagnosis of  heart  failure, cardiomyopathy, or myocardial infarction) provided written  informed consent. Image acquisitions were performed on a 3T Siemens Tim Trio (Siemens Healthcare, Erlangen, Germany)  scanner with a 6-element chest coil and a 24-element spine coil.

Coronal Localizer Image

Image of Respiratory Navigator

(during breath-hold)

Lung

Exhale

Inhale

Detected Diaphragm Location

Liver

Time

Figure 2.2:  Respiratory  navigator gating. (Left) The diaphragm position was measured at the high-contrast interface between the lung (dark) and the liver (bright). (Right) Image of a measured diaphragm position over time during a breath-hold.

2.3.2 Quantification of Inconsistent End-Expiratory Positions

To determine the inconsistency in end-expiratory positions for each subject, a respiratory navigator sequence measured the diaphragm position (Figure 2.2) during

10 consecutive breath-holds. During each breath-hold, the diaphragm position was imaged three times per second over a period of 10 seconds for 30 total measurements. No cardiac image data were collected during these acquisitions. The mode of the

30 diaphragm positions defined the measured end-expiratory  position of that breath- hold. The patient-specific minimum, middle, and maximum end-expiratory positions were defined from the series of 10 breath-holds (Figure 2.3).

2.3.3 DENSE Acquisition

For  each  subject, navigator-gated 2D spiral cine DENSE in  2-chamber and

4-chamber long-axis and basal, mid-ventricular, and apical short-axis orientations of the left ventricle were acquired four times. Specifically, all image orientations were

Measured

Defined Patient-Specific End-Expiratory Positions

Relative

Position (mm)

End-Expiratory Positions

10

Minimum

Middle

Maximum

Exhale

0

-10

d

Inhale

1 Breath-hol  10

Number

Figure 2.3:  A respiratory  navigator was used to measure end-expiratory positions to define the patient-specific minimum,  middle,  and maximum end-expiratory  positions. The minimum position was defined  as being closer to the end-inspiratory position while the maximum position was defined as being closer to the end-expiratory position.

acquired with  the navigator acceptance window prescribed at the patient-specific maximum and minimum end-expiratory positions, and twice in the middle position to quantify variability  in strain independent  of end-expiratory position variability (Figure 2.3).   A  navigator feedback system, which used an angled mirror  and projector screen  placed at  the back of the scanner  bore, was  used to  facilitate quicker acquisitions by enabling subjects to view the navigator acceptance window position in real-time during image acquisition [30].    For each  end-expiratory position, all image orientations  were acquired within a single navigator-gated scan.

Prospective ECG gating was used during DENSE acquisitions. The number of cardiac phases  ranged from  31 to  49 and varied based on subject heart rate. Additional   DENSE  imaging  parameters included:     spiral  interleaves   =  6, FOV = 360x360 mm2, pixel spacing = 2.8x2.8 mm2, slice thickness = 8 mm, TE =

1.1 ms, TR  = 17 ms, variable flip  angle = 20◦,  displacement  encoding = 0.06

cyc/mm  [6],  through-plane dephasing =  0.08 cyc/mm  [31],  CSPAMM  echo suppression [32], and view sharing. A dual-navigator strategy was used, requiring the diaphragm to be within the navigator acceptance window (±3 mm) both before

and after the data acquisition during each R-R interval [33].

2.3.4 DENSE Post-Processing

DENSE  image  data  were  analyzed  using  the   open-source software, DENSEanalysis [34].  For each image orientation, the left ventricular myocardium was   manually  delineated using  epicardial  and  endocardial contours  and  an end-diastolic and end-systolic cardiac phase [35]. Post-processing and segmentation were performed as  described by  Suever et  al.    [35].    Seed  points  indicating unwrapped  phase data were manually selected, and a path-following algorithm was used to unwrap the displacement-encoded  phase data.  The resulting displacement trajectories were further  processed  by applying spatial smoothing and temporal fitting as previously described [36].

Two-dimensional Lagrangian strains were computed from  the  smoothed trajectories over the entire cardiac cycle.  Radial and circumferential strains were computed from the short-axis images and longitudinal strain was computed from the long-axis images. Global peak strains were calculated by averaging the mean strain curves of all the myocardial segments and identifying the peak of the global mean curve. Regional peak strains were computed by averaging the strain curves from all the myocardial segments for a given region and identifying the peak of the regional curve. Segmental peak strains were computed  by identifying the peak of the strain curve for each myocardial  segment.  For peak longitudinal strain computation, pixels within  10% of left ventricular longitudinal length from the most basal and apical regions were excluded  because of the increased noise which is typically observed in the strain curves in those regions. Peak strain was defined  as positive for thickening (radial) and negative for shortening (circumferential and longitudinal).

2.3.5 Statistics

Statistical analyses  were performed using R version  3.2.2 (R  Foundation for Statistical Computing, Vienna, Austria). All continuous  variables were expressed as mean ± standard deviation or range.  Cardiac strains were tested  for normality using a Shapiro-Wilk test.

To quantify mean differences in cardiac strains due to inconsistent end-expiratory positions (minimum, middle, and maximum positions), cardiac strains were compared between the patient-specific  acceptance window positions using a two-way analysis of variance (ANOVA)  with repeated measures with group (healthy vs patient) and acceptance window position as the independent  factors.  A Scheirer-Ray-Hare  test was used for data determined to be non-normally distributed [37]. Using the results of the two-way  ANOVA  or Scheirer-Ray-Hare test, the interaction between group and acceptance window position on cardiac strains was determined. If there was no interaction between group and acceptance window position,  the groups were combined and mean differences due to inconsistent  end-expiratory positions were quantified by comparing cardiac strains between acceptance window positions using a one-way ANOVA with repeated measures with acceptance window  position  as the independent factor. A Friedman test was used for data determined to be non-normally  distributed [37].

To quantify variability due to inconsistent end-expiratory positions, the standard deviations of strains were compared between the inconsistent positions (maximum, middle, and minimum)  and consistent  positions (two  acquisitions at  the middle position) using a Students t-test. For all statistical tests, significance  was defined as p < 0.05. Bland-Altman analysis [38] was used to assess the reproducibility of each measurement using inter-test 95% limits  of agreement defined using the two measurements  from the middle position.  Inconsistency in end-expiratory position across ten separate breath-holds for each subject was reported using both ranges

and standard deviations from the ten breath-holds, and these values were compared between patients and healthy controls.

Power analyses were performed to quantify the ability  of this study to detect meaningful differences  in  strain  between the  different  end-expiratory positions. Because repeated-measures  ANOVAs were used to detect differences, and because equations for  power are not  readily  available for  repeated-measures  ANOVA, simulations were performed to estimate power. Specifically, for each strain, 10,000 iterations were performed.   For  each  iteration,  strain  values for  the minimum end-expiratory position were randomly drawn from a normal distribution using the mean and standard deviation across  the subjects measured  in  this  study.   The number of strain values drawn corresponded with the number of subjects (healthy and patients combined). For a given difference to detect, δ, values at the two other end-expiratory positions were calculated by adding δ/2 and δ  to the values at the minimum  position.     Measurement   variability   was   then  added to  those two end-expiratory positions by drawing random values from a normal distribution with zero mean and a standard deviation equal to  the  measured average standard deviation of the differences  between any two  positions.   In  this  manner, each iteration simulated a mean difference  of δ  between the minimum and maximum breath-hold positions and included typical  inter-test measurement variability.   The percentage of iterations for which a repeated-measures ANOVA yielded a significant result (p < 0.05) was the estimate of power. The 95% confidence interval of that estimate was calculated from the normal approximation to the binomial distribution with N = 10,000. The values of δ  that yielded at least 80% power were reported separately for global and regional strains.

2.4  Results

was excluded due to movement during imaging,  so data from the remaining 9 healthy subjects are reported. Data from the basal and apical DENSE images of these subjects were previously  used to assess the variability of left ventricular torsion [39].

2.4.1 Inconsistent End-Expiratory Positions

As previously reported [39], the average range of end-expiratory positions were not significantly different (p = 0.94) between the healthy (10.1 ± 4.8 mm) and patient (10.3 ± 4.2 mm) groups (total range of 4-19 mm). Since range is sensitive to outliers, the standard deviation of end-expiratory position was also compared between groups, and similarly there were no significant differences (3.1 ± 1.3 mm vs 3.4 ± 1.7 mm, p = 0.70) [39].

2.4.2 Differences and Variability in Peak Strains

There was  no interaction between group (healthy vs patient)  and navigator acceptance window position for peak strains (Table 2.1), thus the remaining analyses were performed with all subjects combined. Neither global, regional, nor segmental peak strains  were significantly different  as  a function  of  acceptance window position (Table 2.1; Tables A.1 and  A.2 in Supplemental). Moreover, the differences in mean strain between any two acceptance window positions  were each smaller than their corresponding inter-test 95% limits of agreement (Table 2.1). For example, mean global circumferential strain across  acceptance  window positions ranged from -16% to  -17%; the difference was  1%, which is smaller than  the corresponding inter-test 95% limits of agreement of ±1.7% (Table 2.1). Finally, the standard deviations in peak strains were not significantly different between inconsistent (minimum, middle, and maximum) and consistent (repeated measurements  at  middle position) acceptance  window positions for  all  subjects combined (Table 2.2). With at least 80% power, this study had the ability to detect strain  differences of  4.7%, 1.0%, and  1.7% (absolute) between  end-expiratory

Table 2.1:  Global  and regional peak strains (mean   standard deviation) from the three acceptance window positions (minimum, middle,  and maximum)  for all subjects combined.

Measurement

Radial Strain (%)

Acceptance Window  Position

Minimum 	Middle 	Maximum

p-value†	95% LoA 	p-value‡

Global	29 ± 12	29 ± 12	30 ± 13	0.95	±7.9	0.99

Base	37 ± 15	35 ± 14	37 ± 14	0.95	±13.1	0.89

Mid-Ventricle 	28 ± 11	29 ± 14	28 ± 13	0.77	±10.4	0.51

Apex	26 ± 12	27 ± 10	29 ± 16	0.78	±15.8	0.79

Circum. Strain (%)

Global	-16 ± 4	-17 ± 4	-17 ± 5	0.57	±1.7	0.65

Base	-15 ± 4	-15 ± 4	-15 ± 4	0.83	±3.6	0.71

Mid-Ventricle 	-16 ± 4	-17 ± 4	-17 ± 4	0.17	±2.1	0.78

Apex	-19 ± 5	-19 ± 5	-19 ± 5	0.98	±4.6	0.93

Long. Strain (%)

Global	-12 ± 4	-12 ± 3	-13 ± 4	0.44	±3.2	0.48

2ch	-13 ± 3	-12 ± 3	-13 ± 4	0.94	±6.2	0.75

4ch	-12 ± 4	-13 ± 4	-13 ± 4	0.84	±4.1	0.38

†Results from test comparing acceptance window positions

‡Results from test comparing interaction between group (healthy vs patient) and acceptance

window position

positions for  global radial,  circumferential, and longitudinal  strain,  respectively

(Table 2.3). Additionally, this study had at least 80% power to detect differences of

8.9%, 2.2% and 2.6% for regional radial, circumferential, and longitudinal strain, respectively (Table 2.3).

2.5  Discussion

Quantification of cardiac strains typically requires a series of image acquisitions performed during end-expiratory breath-holds. This study explored the effects of inconsistent end-expiratory positions on the quantification of left ventricular cardiac strains.    The  results of  the  study  showed that  1)  inconsistent  end-expiratory positions had minimal  effect on the quantification of global and regional peak strains compared to inter-test variability  for a given imaging location; and 2) the

Table 2.2:  Standard  deviation of global and regional peak strains across inconsistent (maximum,  middle, minimum)  and consistent (middle and repeated  middle)  acceptance window  positions.   Values reported as  mean standard deviation.

Measurement	Acceptance Window  Positioning 	p-value

Inconsistent	Consistent

Radial Strain (%)

Global	3.2 ± 1.7	2.3 ± 1.8	0.17

Base	4.9 ± 2.5	3.5 ± 3.3	0.18

Mid-Ventricle 	4.3 ± 2.5	2.5 ± 2.8	0.10

Apex	5.7 ± 4.1	4.3 ± 3.9	0.16

Circumferential Strain (%)

Global	0.7 ± 0.4	0.5 ± 0.3	0.10

Base	1.0 ± 0.5	1.1 ± 0.7	0.41

Mid-Ventricle 	0.7 ± 0.3	0.6 ± 0.5	0.55

Apex	1.4 ± 0.9	1.4 ± 0.9	0.95

Longitudinal Strain (%)

Global	1.1 ± 0.7	0.8 ± 0.9	0.27

2ch	1.4 ± 1.1	1.5 ± 1.6	0.89

4ch	1.4 ± 1.6	1.2 ± 0.9	0.72

Table 2.3: Power analyses for studys ability to detect a difference in global and regional strain between different end-expiratory  positions.

Measurement	Power (%) 	Difference To Detect

(absolute, %)

Radial Strain (%)

Global	80.9 ± 0.8	4.7

Base	96.1 ± 0.4

Mid-Ventricle 	98.4 ± 0.2

Apex	80.2 ± 0.8

Circumferential Strain (%)

8.9

Global	80.5 ± 0.8	1.0

Base	99.6 ± 0.1

Mid-Ventricle 	100 ± 0.0

Apex	80.6 ± 0.8

Longitudinal Strain (%)

2.2

Global	80.0 ± 0.8	1.7

2ch	95.0 ± 0.4	2.6

4ch	80.5 ± 0.8

variability  of global and regional peak strains was similar between inconsistent and consistent  end-expiratory positions. Importantly,  these findings provide assurance that  the  measurement  of  cardiac strains is  relatively  robust  with  respect to inconsistent end-expiratory positions.

Peak strains vary throughout the left ventricle. For example, we found that the magnitude of circumferential strain was 2% (absolute) higher in the apical region than the base–in agreement with previous studies [14, 15, 16, 17, 18, 19]–and radial strain was 9% (absolute) higher in the basal region than the apex.  Due to these strain gradients,  we hypothesized that the displacement of the heart due to motion of the diaphragm with  respect to the imaging plane would create differences  in measured strains.    For  example,   we  might  expect that  radial  strains for  the maximum end-expiratory position (i.e.,  maximal exhalation) would be  lower in magnitude compared to  the minimum end-expiratory position due to  the heart being imaged more apically.   We  also might  expect this  to  manifest as  higher variability in strains across different end-expiratory positions compared to consistent end-expiratory positions.   The likely explanation for finding that there is no difference in  strains between end-expiratory positions is  that,   because  the longitudinal axis of the heart (base to apex) is not necessarily perpendicular  to the diaphragm plane, a 10 mm translation in the diaphragm position does not directly correspond with a 10 mm translation of the heart through the imaging plane.

Previous studies suggest that regions of the heart could displace at least 3 and possibly up to 14 mm through the fixed imaging plane between breath-holds [26,

27, 28, 29]. Our study had an average range of end-expiratory diaphragm position between breath-holds of approximately 10 mm, which is consistent  with  previous studies. Since the imaging slice thickness is 8 mm, even with a 14 mm through-plane displacement, there is likely not much difference in the acquired data from the imaged heart locations compared to the imaging plane location. Overall, since there were no

significant differences in peak global, regional, and segmental strains between end- expiratory positions, patient end-expiratory diaphragm position does not have to be monitored when performing breath-hold DENSE acquisition for single image analyses.

This  studys  goal was  to  quantify  the  effects of  inconsistent  end-expiratory positions on cardiac strains by computing the differences in strain between different end-expiratory positions. Thus, it was important for this study to detect meaningful strain differences between  different  patient-specific end-expiratory positions.  This study had at least 80% power to detect global strain differences of 4.7%, 1.0%, and

1.7% and regional strain differences  of 8.9%, 2.2%, and 2.6%, between different end-expiratory positions for radial,  circumferential, and longitudinal strain, respectively. Importantly, the studys detectable  difference was similar to or smaller than previously reported values of inter-test limits of agreement for circumferential strain and radial strain [6].   Notably,  in  some regions, the power to  detect a meaningful difference  was much higher (close to 100%) indicating that,  in those regions, this study may have had the ability  to detect even smaller than reported detectable differences.

We used DENSE to investigate our hypothesis that a patients normal variability in  end-expiratory position  between image  acquisitions significantly  affects the quantification of cardiac strains. DENSE was chosen to test our hypothesis because it has been previously shown to have good reproducibility [5], can be acquired in high spatial resolution [4, 13], and enables straightforward computation of cardiac strains. However, our findings should generalize to other image acquisitions that are used to derive  measures of cardiac strains such as echocardiography,  tagged MRI, etc.

We  used respiratory navigator gating to acquire the DENSE cardiac images, which  reduces respiratory artifacts  during  image acquisition, so  we  could not measure the effect of inconsistent  end-expiratory position during breath-holds on

the derived strains. It would be beneficial to quantify the amount of end-expiratory position variability during breath-hold cardiac MR image acquisition and determine whether the  magnitude of  inconsistent  end-expiratory positions correlates with changes  in strain values.  An example would be  to explore whether inconsistent end-expiratory positions during a breath-hold DENSE scan causes blurring due to motion and results in lower strain magnitudes.

This study examined the effects of inconsistent end-expiratory positions on cardiac strains in a small patient sample.  It would be beneficial  to investigate this effect in a larger patient sample who have heterogeneous contraction patterns, for example, due to post-myocardial infarction.   These patients may have  steeper gradients in strain across infarcted to non-infarcted tissue regions [40].  Therefore, we cannot definitively say that the effects of inconsistent end-expiratory positions in that setting are similarly small and negligible. Future studies should investigate strain variability due to inconsistent end-expiratory positions in patients who have infarcted tissue in specific regions (e.g. anterior vs inferior).

In  conclusion, the  quantification  of  peak left  ventricular  cardiac strains is relatively insensitive to normal variations in end-expiratory  positions between image acquisitions. Since there were no differences in peak strain between end-expiratory positions, patient end-expiratory diaphragm position does not have to be monitored when performing breath-hold DENSE acquisition for single image analyses. These findings should generalize  to  other image acquisitions that  are used to  derive measures of cardiac strains.

CHAPTER 3

USING A RESPIRATORY NAVIGATOR  REDUCES VARIABILITY WHEN QUANTIFYING LEFT VENTRICULAR  TORSION

Adapted from Hamlet SM, Haggerty CM, Suever JD, Wehner GJ, Andres KN, Powell DK, Charnigo RJ, Fornwalt BK. Using a Respiratory Navigator Significantly Reduces Variability when Quantifying Left Ventricular Torsion with Cardiovascular Magnetic Resonance. Journal of Cardiovascular Magnetic Resonance. 2017. (in Press)

The purpose of this work was  to determine the effects of using a respiratory navigator on the variability  of left  ventricular  torsion derived from  spiral cine displacement  encoding with stimulated echoes (DENSE) MRI. In this chapter, we discuss the two separate experimental  protocols (using 1.  enforced  and 2.  natural variability   in  end-expiratory position)  used to  test  the  hypothesis that  high inter-test  variability   in  left  ventricular  torsion  is  partly   due  to  inconsistent breath-hold positions during serial image acquisitions, which could be significantly improved by using a respiratory navigator for cardiac MRI based quantification of left ventricular torsion.

3.1  Synopsis

Background:  Left ventricular (LV)  torsion is an important  indicator of cardiac function that is limited by high inter-test variability  (50% of the mean value). We hypothesized that this high inter-test variability is partly due to inconsistent breath- hold positions during serial image acquisitions, which could be significantly improved by using a respiratory navigator for cardiovascular  magnetic resonance (CMR) based quantification of LV torsion.

Methods:  We assessed respiratory related variability in measured LV torsion with

two  distinct experimental protocols. First,  17 volunteers were recruited for CMR with  cine displacement  encoding with  stimulated echoes  (DENSE)  in  which a respiratory  navigator  was   used to  measure and  then  enforce variability   in end-expiratory position between all LV basal and apical acquisitions. From these data, we quantified the inter-test variability  of torsion in the absence and presence of enforced end-expiratory position variability,  which established an upper bound for the expected torsion variability.  For the second experiment (in 20 new, healthy volunteers), 10 pairs of cine DENSE basal and apical images were each acquired from consecutive breath-holds  and consecutive navigator-gated  scans (with a single acceptance position).   Inter  test variability of torsion was compared  between  the breath-hold and navigator-gated scans  to quantify the variability  due to natural breath-hold variation.     To  demonstrate the  importance of  these variability reductions,  we quantified the reduction in sample size required  to detect a clinically meaningful  change in LV torsion with the use of a respiratory navigator.

Results: The mean torsion was 3.4±0.2 ◦/cm.  From the first experiment, enforced

variability   in  end-expiratory position  translated  to  considerable variability   in measured torsion (0.56±0.34 ◦/cm),  whereas  inter-test variability  with  consistent

end-expiratory position was 57% lower (0.24±0.16 ◦/cm, p>0.001).  From the second

experiment, natural respiratory variability from consecutive breath-holds  translated to a variability in torsion of 0.24±0.10 ◦/cm, which was significantly higher than the

variability  from  navigator-gated scans (0.18±0.06  ◦/cm,  p=0.02).    By  using a

respiratory navigator with DENSE, theoretical sample sizes were reduced from 66 to

16 and 26 to 15 as calculated  from the two experiments.

Conclusions: A substantial portion (22−57%) of the inter-test variability  of LV torsion can be reduced by using a respiratory navigator to ensure a consistent breath- hold position between image acquisitions.

Keywords: Left Ventricular Torsion, Breath holds, DENSE, Respiratory Navigator

Gating, cardiovascular magnetic resonance

3.2  Background

Left ventricular (LV) torsion is an important indicator of cardiac function [41, 42]; however, the quantification of torsion is limited by poor inter-test reproducibility. For example, a previous study with myocardial tagging demonstrated that the inter-test variability of torsion represented nearly 50% of the mean value [19]. This substantial variability reduces prognostic value for individual patients and leads to larger required sample sizes for research studies to detect meaningful differences or changes. Previous studies have reported that sample sizes ranging  from 80 107 are required to detect a

10% relative difference in torsion with 90% power [19, 43, 44]. Reducing variability and lowering  required  sample sizes is important to improve the clinical and research utility  of torsion.

LV torsion is typically quantified as the gradient of twist along the longitudinal axis of the heart. This gradient is computed using twist derived from two short axis images (basal and apical) of the LV  and the longitudinal distance between  the images [19] (Figure  3.1).    End-expiratory  breath-holds are used to  minimize respiratory motion artifacts, and the basal and apical short axis images are typically acquired during separate breath-holds.  When post processing  the image data to compute LV  torsion, the longitudinal distance between the short axis images is calculated from  either  A)  assumptions derived from  an additional  longitudinal image (echocardiography)  or B) information specifying the location of the imaging planes in  3D  space taken  from  the  Digital  Imaging and  Communications in Medicine (DICOM)  image header (cardiovascular magnetic resonance  [CMR]). A confounding factor that is not considered is that the exact end-expiratory position may differ by up to 13 mm between separate breath-holds  [21, 22, 23, 24, 25], which creates differences in heart position between the basal and apical image acquisitions (Figure 3.2). We hypothesized that inconsistent end-expiratory diaphragm positions during serial breath-holds accounts for a significant  portion of the variability  in

Base

Base

Apex

ϕb

d

Apex 	ϕa

ϕ = ϕa  - ϕb

ϕ

= 	d

Figure  3.1:   Computation  of LV torsion from  basal  and  apical images. The curved arrows represent  the relative  twist  along the longitudinal axis of the left ventricle. LV twist (φ) was measured  as the difference in rotation between the apex (φa) and base (φb) (twist direction shown as viewed from foot to head). Torsion τ  was computed  as LV twist divided by the distance (d) between basal and apical image locations.

measured LV torsion and that this variability could be reduced by using CMR based quantification of LV torsion with a respiratory navigator.

3.3  Methods

Respiratory related variability  in measured LV  torsion was  assessed with  two distinct  experimental protocols:  1)  using enforced  variability  in  end-expiratory position   between  acquisitions  and   2)   allowing   for   natural   variability    in end-expiratory position between acquisitions.  The former experiment was performed to establish an upper bound on respiratory related variability  in torsion, while the latter mimics a more relevant  clinical setting.  In both experiments,  the effect of using a respiratory navigator to  ensure a consistent  end-expiratory position on torsion variability  was  also quantified.    The  local Institutional   Review Board approved the study protocols, and all subjects provided written informed consent.

Cardiac Imaging Plane

Original Diaphragm Position

Current Diaphragm Position

End-Expiration	Mid-Inspiration	End-Inspiration

Figure  3.2:  Real  time  images of the  diaphragm  as it  translates during a respiratory  cycle.  During respiration, diaphragm motion causes the heart to translate a substantial distance through the fixed imaging plane.

3.3.1 LV Motion Quantification

Imaging was performed on a 3T Siemens Tim Trio (Siemens Healthcare, Erlangen, Germany) with a 6 element  chest coil and a 24 element  spine coil.  LV twist  was measured at basal and apical short axis locations in both experiments using 2D spiral cine Displacement Encoding with Stimulated Echoes (DENSE) CMR [45, 46].

The basal and apical short axis locations were defined as  follows:  On a four chamber image, five short axis slices were planned equidistant across the end systolic endocardial ventricular long axis length.   The slices were planned such that  the outermost  slices did not extend beyond the mitral valve plane and endocardial apex, respectively. The second and fourth slices of this stack were defined as the basal and apical short axis locations.   Imaging parameters were:  spiral interleaves  = 6, interleaves per frame = 2, FOV = 360x360 mm2, pixel spacing = 2.8x2.8 mm2, slice thickness = 8 mm, TE/TR = 1.1/17 ms, temporal resolution = 34 ms, variable flip angle = 20◦, displacement encoding = 0.06 cyc/mm [6], through plane dephasing =

0.08 cyc/mm [31], CSPAMM echo suppression  [32], view sharing, prospective ECG

gating, and a respiratory navigator with an acceptance window of ±3 mm.

DENSEanalysis [34] was  used to  derive  LV  twist  from the DENSE images. Epicardial and endocardial contours were manually delineated on the  DENSE magnitude images at  end-diastolic and end-systolic cardiac phases  [35].    Post processing  was  performed as  previously described  [35].   A  semi-automatic path following algorithm was used to unwrap the displacement encoded phase data. The resulting displacement trajectories were further  processed  by applying spatial smoothing and temporal fitting [36].

LV  twist  was  computed over the cardiac cycle relative  to  the centroid of the endocardial boundary at end-diastole. The distance between the basal and apical image locations was calculated from the DICOM headers. LV torsion was computed as the difference in rotation between the apex and base (φ) divided by the distance (d) between the basal and apical image locations [19, 47, 48] (Figure 3.1).

3.3.2 Experiment 1: Enforced End-Expiratory Variability

Ten healthy volunteers with no known cardiovascular  disease or chronic illnesses and seven patients with a history of heart disease (known diagnosis of heart failure, cardiomyopathy, or myocardial infarction) were recruited.  We  first quantified the end-expiratory variability  for each subject by acquiring respiratory navigator measurements  (90180 cross pair configuration; Figure 3.3) of 10 consecutive,  10 second breath-holds.  No cardiac image data were acquired, but the mode position of each  breath-hold was  retained to identify  subject specific minimum, middle and maximum end-expiratory positions of the diaphragm across  the 10 breath-holds (Figure 3.4). These subject specific positions were then used to define the locations of the  navigator acceptance  windows for  subsequent  acquisitions of respiratory navigator-gated DENSE. Specifically, the basal and apical slices were both acquired with the navigator acceptance window at each of the three positions. Moreover, the acquisitions at  the middle acceptance  window location were repeated to  define inter-test  variability   when  ensuring a  consistent   position  with   a  respiratory

Coronal Localizer

Image of Respiratory Navigator

Breath-hold 1 	Breath-hold 10

Exhale

Lung

Liver

Detected Diaphragm Location

11mm

Inhale

Time

Figure 3.3: Respiratory navigator gating. (Left) The diaphragm position is at the high contrast interface between the lung (dark) and the liver (bright).  (Right) Image of a measured diaphragm position over time for separate breath-holds.  For this subject, there was an 11 mm difference in end-expiratory position between breath- hold 1 and 10.

navigator. For all scans, the image of the respiratory navigator was projected to the subjects in real time during DENSE acquisition, which helped to ensure consistent efficiency [33] across scans despite varying acceptance locations.

With  three independent measurements at both LV locations, nine permutations of torsion were calculated from the possible combinations (Figure 3.5), providing an estimate of torsion variability due to inconsistent end-expiratory positions. This variability in torsion was compared to the inter-test variability (i.e., comparing the two torsion measures acquired  at the middle navigator acceptance position, Figure 3.5) to isolate respiratory position effects.

3.3.3 Experiment 2: Natural End-Expiratory Variability

We  next sought  to  quantify  the effects of natural  end-expiratory variability. Twenty  new healthy volunteers were recruited.   In  these subjects, 10 basal and apical images were each  acquired with  two  protocols:   1)  during  consecutive

Measured End-Expiratory

Diaphragm Positions

60

Defined Subject-Specific End-Expiratory Positions

Maximum 	Middle 	Minimum

Exhale

Relative Position (mm)

50

40

30

20

Inhale

10

0

1	10

Breath-hold Number

Figure  3.4:  Measured  end-expiratory  diaphragm  positions were used to define subject  specific maximum,  middle,  and minimum  end-expiratory positions.  The maximum diaphragm position was defined  as being  closer to the end-expiratory position while the minimum diaphragm position was defined as being closer to the end-inspiratory position.

Torsion Permutations

1 	2 	3 	4 	5 	6 	7 	8

For Inter-test

Comparison

9 	Repeat 5

Base

Apex

Maximum Maximum Maximum

Maximum    Middle 	Minimum

Middle 	Middle 	Middle

Maximum    Middle 	Minimum

Minimum  Minimum

Maximum    Middle

Minimum

Minimum

Middle

Middle

Figure 3.5: The nine possible torsion permutations were constructed from three basal and three apical images. One basal and apical image was acquired for each subject specific end-expiratory  position (maximum, middle, and minimum). Image acquisitions were repeated at the middle position to assess inter-test variability (far right).

breath-holds, and 2) during consecutive navigator-gated  acquisitions  with a single acceptance window location. In each case, the 10 image pairs were used to derive 20 measurements of LV torsion, by combining each basal twist measurement with the two  closest apical twist  measurements  in  the temporal sequence.  The torsion variability between these protocols was then quantified to compare the differences  as a    result    of    consistent    (navigator-gated)   and   inconsistent    (breath-hold) end-expiratory positions. Importantly, to monitor the end-expiratory position of the breath-hold acquisitions,  the  scans  were acquired with  the respiratory navigator

enabled, but with a wide (±50 mm) acceptance window width that never resulted in

the exclusion of acquired image data.

3.3.4 Statistics

Statistical analyses  were performed using R version  3.2.2 (R  Foundation for Statistical Computing, Vienna, Austria). All continuous  variables were expressed as mean ± standard deviation and group means were compared using Student’s t tests. Pearson correlation was used to observe associations between continuous  variables.

For experiment 1, the inter-test variability  of torsion was quantified using 95% inter-test limits of agreement of the two middle navigator acceptance window scans. To test for an overall difference in variability between inconsistent and consistent end- expiratory positions, the LV torsion permutations from the variable end-expiratory positions  were compared to the 95% inter-test limits of agreement using a binomial test to evaluate whether values fell within the 95% limits significantly less than 95% of the time. The root mean squared error (RMSE) was then computed to quantify the differences in variability. Specifically, the RMSE for the consistent end-expiratory position was computed by computing the mean squared error (MSE) of the two middle acceptance window scans and taking the square root. The RMSE for the LV torsion permutations was computed  by separately computing the MSE of the permutations with respect to each of the two middle acceptance window  scans, averaging the MSEs,

and taking the square root.

For experiment  2, breath-hold and navigator-gated  acquisitions were compared by  computing the  standard deviations of  the  20 respective  measurements  and performing a Student’s t test.  Variability in torsion was also quantified using 95% inter-test limits of agreement, which were computed using the standard deviation of the difference between consecutive pairs of torsion measurements.  For all statistical tests, significance was defined as p < 0.05.

3.3.5 Theoretical Sample Size Calculation

To quantify the effects of the differences in torsion measurement variability,  we computed theoretical sample  sizes required to detect a clinically meaningful change in LV torsion for each experimental  condition. Study sample sizes required  to detect a 10% relative difference in LV torsion with a power of 90% and a significance level of 0.05 were computed using the standard deviation of the inter-test differences in torsion (α) and the equation:

2

·	·

n = f (α, P )  σ2

δ2

(3.1)

where n is the sample size per group, α is the significance level, P is the power, f

is the value of the factor for different values of α and P (f = 10.5 for α = 0.05 and P

= 0.90), and δ is the magnitude of the difference to be detected [49]. To determine the improvement in sample size compared to other modalities, sample sizes calculated  by this formula were compared to those calculated  based on data from previous studies that quantified LV torsion.

3.4  Results

For experiment 1, ten healthy volunteers (Age: 22 ± 6 years, Range: 19−38 years,

60% female) and seven patients  (Age: 57 ± 8 years, Range: 45−67 years, 43% female)

were enrolled. One healthy volunteer was excluded due to movement during imaging, so data from the remaining nine healthy volunteers are reported. For experiment 2,

20 healthy volunteers (Age: 25 ± 4 years, Range: 20−34 years, 60% female) were

enrolled.

3.4.1 Inconsistent End-Expiratory Positions

From  experiment   1,   the  intra-subject  range  and  standard  deviation  of end-expiratory  positions were 10.2 ± 4.4 mm and 3.3 ± 1.4 mm, respectively. There was no significant  difference in the range or standard deviation of end-expiratory position between healthy and patient groups (p = 0.94 and p = 0.70, respectively; Figure 3.6). From experiment 2, the intra-subject range and standard deviation of end-expiratory positions over 20 breath-holds were 13.9 ± 10.5 mm and 3.8 ± 3.1 mm, respectively.

3.4.2 Torsion

DENSE images and displacements from a representative subject show the relative twist differences between the base and apex at end-systole (Figure 3.7). Table 3.1 summarizes the LV torsion results for each protocol. From experiment 1, the inter- test limits of agreement at a consistent position were ±0.6 ◦/cm, and the binomial test indicated that the variability  in LV torsion due to enforced variability  in end- expiratory position was significantly higher than the variability at a consistent end- expiratory position (p < 0.001). Specifically, the RMSE of LV torsion permutations across end-expiratory positions was 0.56 ± 0.24 ◦/cm (range: 0.2−1.3 ◦/cm), while the RMSE from a consistent end-expiratory position was 57% lower (0.24 ± 0.16 ◦/cm). Moreover, there was a moderate correlation  across subjects between the torsion RMSE and the range of end-expiratory positions (r = 0.50, p = 0.049, Figure 3.8). Finally, the mean LV  torsion for consistent  end-expiratory positions was  not significantly different between the healthy (3.6 ± 1.2 ◦/cm) and patient (3.2 ± 1.3 ◦/cm) groups

Range of End-Expiratory Position (mm)

a 25

20

15

p = 0.94

b	7

p = 0.70

Standard Deviation of

End-Expiratory  Position (mm)

6

5

4

10	3

2

5

1

0

Healthy	Patient

0

Healthy	Patient

Figure  3.6:  Inconsistent  end-expiratory  positions across ten consecutive breath  holds in  patients  and  healthy  controls.   There were no significant differences in either the range (a) or standard deviation (b) of end-expiratory position between the healthy and patient groups. Solid red lines denote the mean for each group.

Magnitude 	X-Displacement 	Y-Displacement 	2D Displacements

Base

Left 	Right     Up 	Down

Apex

Left 	Right     Up 	Down

Figure 3.7: DENSE images from a representative subject show the relative twist differences between the basal and apical images at end systole. Twist in the basal region is predominantly in the clockwise direction, while the apex is predominantly counter clockwise.

(p = 0.30).

For  experiment   2,  consecutive   breath-holds yielded  a  significantly  larger standard deviation of LV torsion compared to consecutive navigator scans (0.24 ±

0.10 ◦/cm vs 0.18 ± 0.06 ◦/cm, p = 0.02). There was a moderate correlation across

subjects between the standard deviation of torsion and the standard deviation of end-expiratory position (r  = 0.34, p  = 0.03, Figure 3.9).   The 95% limits  of agreement from the consecutive breath-hold scans and consecutive navigator scans

were ±0.74 ◦/cm and ±0.56 ◦/cm, respectively.

3.4.3 Theoretical Sample Sizes

The theoretical sample sizes required  to detect a 10% relative difference in peak torsion (δ  = 0.34 ◦/cm)  from each experimental protocol are shown in Table 3.2. From  both  experiments, using a  respiratory navigator  with  DENSE produced similar sample size estimates  (n = 16 and 15). By comparison,  sample sizes based

Table 3.1:  Mean  (± standard deviation)  of torsion across the volunteers

within  each experiment.

Method  (experiment) 	Torsion (◦/cm) 	p-value

Experiment 1*

Enforced inconsistent positions	3.4 ± 0.4	0.85

Consistent positions with navigator	3.4 ± 0.2

Experiment 2

Enforced inconsistent positions	3.6 ± 0.3	0.32

Consistent positions with navigator	3.5 ± 0.2

*Reported values are from combined group of healthy and patient volunteers

RMSE of Torsion (°/cm)

1.5

1

Healthy

Patient

r  = 0.50

p = 0.049

0.5

0

0	5	10	15	20

Range of End-Expiratory Position (mm)

Figure 3.8:   Variability of torsion due to enforced inconsistent end- expiratory  positions versus the  subject specific range of end-expiratory position. There was a moderate positive correlation between RMSE of LV torsion due to inconsistent expiratory positions and the range of end-expiratory position (r

= 0.50, p = 0.049). The dashed gray line illustrates the linear best fit.

Standard Deviation

of Torsion (°/cm)

0.5

0.4

0.3

0.2

0.1

0

Navigated

Breath-hold

r  = 0.34

p = 0.03

0	5	10	15

Standard Deviation of

End-Expiratory Position (mm)

Figure 3.9:  Variability of torsion due to naturally  inconsistent end- expiratory  positions versus the standard deviation of end-expiratory position. There was a moderate positive correlation between the standard deviation of LV torsion and the standard deviation of end-expiratory position (r = 0.34, p =

0.03). The dashed gray line illustrates the linear best fit.

on measurements  with  variable end-expiratory positions were up to  313% higher. Additionally,  compared to  other modalities, using a respiratory navigator with DENSE provided a 80 to 86% reduction in the required sample size compared to CMR tagging [19], CMR feature tracking [43], and 3D speckle  tracking echocardiography  [44] (Table 3.2).

3.5  Discussion

This study explored the effects of inconsistent end-expiratory diaphragm positions on the quantification of LV torsion and how enforcing a consistent  end-expiratory position with a respiratory navigator can significantly reduce inter-test variability of measured LV torsion. Our primary findings include 1) using a respiratory navigator with DENSE to enforce a consistent end-expiratory position reduced the variability in measuring torsion by 2257%; 2) this decreased variability reduced the required sample sizes to detect a 10% relative  difference in torsion from n = 66 to n = 16 (from enforced variability  to consistent) and n = 26 to n = 15 (from natural variability to consistent); 3) the variability  of LV  torsion due to inconsistent  end-expiratory

Table 3.2:  Sample sizes required to detect a 10% relative  change in LV

torsion calculated using data in this and previous studies.

Method  (experiment) 	Sample Size (n)

Experiment 1

Enforced inconsistent positions	66

Consistent positions with navigator	16

Experiment 2

Enforced inconsistent positions	26

Consistent positions with navigator	15

Previous Studies

| CMR Tagging [19] | 107 |
| --- | --- |
| CMR Feature Tracking [43] | 81 |
| 3D Speckle Tracking [44] | 80 |

positions had a modest correlation with the variability  in end-expiratory positions, such that greater inconsistency in end-expiratory positions was associated with larger errors in measured LV torsion. Regarding inconsistency in end-expiratory positions,

within each subject, substantial inconsistency existed with a mean range of 10 ± 4

mm and 14 ± 10 mm in experiment 1 and 2, respectively, which was similar to that reported previously (7 to 13 mm) [21, 22, 23, 24, 25].

LV torsion is an important indicator of cardiac function because it integrates the three dimensional deformation of the complex myocardial fiber architecture into a single metric [41, 42]. In many disease states, small disruptions in normal cardiac geometryand thus  torsionmay  precede appreciable changes in  global  cardiac function.   For example, previous studies in mice and canines have  reported that changes in torsion precede changes in ejection fraction and volumes in obese animals compared to healthy controls [50, 51]. Previous human studies have reported that LV  torsion differs between younger and older populations, and is also reduced in patients  with   hypertrophic  cardiomyopathy,  valvular  heart  disease, previous myocardial infarction,  and dilated cardiomyopathy compared to healthy controls [42, 47, 52, 16, 53, 54]. Therefore, accurate and reproducible quantification of LV

torsion may provide a robust,  clinically  relevant  marker of cardiac health and function.

For LV torsion to be a useful clinical measurement, minimizing the magnitudes and sources of measurement  error is important.   A previous CMR tagging study reported mean torsion values of 3.4 ◦/cm with inter study 95% limits of agreement

of ±1.6 ◦/cm,  representing a large percentage of the mean [19].  Previous CMR

feature tracking studies have reported inter-test limits of agreement  of ±0.9 ◦/cm

[43, 55].   Using DENSE CMR, we  observed  a similar mean of 3.4 ◦/cm  for all subjects combined in  experiment  1 and 3.5 ◦/cm  in  experiment  2, and smaller inter-test 95% limits  of agreement  from the breath hold scans  in  experiment  2

(±0.74 ◦/cm).    However,  the observed  inter-test  95% limits  of agreement  were

considerably smaller when using a respiratory navigator (±0.6 ◦/cm  and ±0.56

◦/cm for experiments 1 and 2, respectively). An important distinction between the present study and previous studies, apart from CMR sequence differences,  is control of the end-expiratory position when quantifying the inter-test variability.

From experiment 1, by comparing the variability of LV torsion inclusive of enforced inconsistent end-expiratory positions (0.56 ± 0.34 ◦/cm) to the variability without this

inconsistency (0.24 ± 0.16 ◦/cm), we determined  that using a respiratory navigator

to ensure a consistent end-expiratory position reduced the variability in measured LV torsion by 57%. In experiment 2, using a respiratory navigator reduced the variability in measured LV torsion by 22% compared to the variability in LV torsion inclusive of naturally inconsistent end-expiratory positions.

In this study, we examined variability in measured LV torsion. A previous CMR study examined the bias in LV twist and circumferential longitudinal (CL) shear angle between different acquisition techniques, including breath holds and free breathing [56]. In agreement with that previous study [56], we did not observe a bias in torsion between breath hold and navigator gated scans (Table 3.1).

To detect a 10% relative difference in peak LV torsion, experiment 1 found that using DENSE with a respiratory navigator required a sample  size of only n = 16 subjects, which is about 76% lower than the sample size required when using DENSE without a respiratory navigator (n = 66). In experiment 2, we found similar results where using DENSE with natural respiratory variability required a sample size of 26 compared to using DENSE with a respiratory navigator (n = 15). Using a respiratory navigator with DENSE provided a 80 to 86% reduction in the required sample size compared to CMR tagging [19], CMR feature tracking [43], and 3D speckle tracking echocardiography [44].

These findings have  meaningful implications for  future  CMR  based quantification of LV torsion in the clinical and research settings. First, acquisition of LV torsion data using a respiratory navigator should be employed, where feasible, to minimize variability.    This approach is not typical  in  the majority  of published papers reporting torsion and may reduce clinical feasibility of such data acquisition; however, the additional effort appears  justified by the considerable  reduction in variance. If inconsistency in end-expiratory position is not addressed with the data acquisition, then it is important to incorporate  effects of inconsistent end-expiratory position into the assessment  of the standard error of measurement  for LV torsion, which will  substantially increase  needed sample  sizes for research trials  or reduce prognostic value for individual subjects.

These results also have  important  implications for echocardiography.    While operators may be  able to correct for inconsistency in end-expiratory position by adjusting the position of the probe, it is unlikely that the operator can recreate the exact distance between each short axis image that was measured from the long axis image.  Because inconsistent  end-expiratory positions is a source of measurement variability  in measured LV torsion in CMR, the discrepancy in distances may be a source of substantial variability in measured LV torsion in echocardiography.

We used spiral cine DENSE to investigate our hypothesis that inconsistent end- expiratory positions accounts for a significant portion of the variability in measured torsion and that inter-test reproducibility could be improved by using a respiratory navigator. We chose to use spiral cine DENSE to investigate our hypothesis since it allows for simple quantification of mechanics,  has good spatial resolution, has good reproducibility, and includes a respiratory navigator, which allows control of the end- expiratory position during image acquisition [45, 5, 4, 13].  However, our findings should generalize to all other imaging modalities that use short axis images to quantify torsion.

3.5.1 Limitations and Future Directions

We examined the effects of variable end-expiratory position on LV torsion in a small patient sample. It may be beneficial to examine these results in a larger, more heterogeneous patient sample to determine whether specific diseases affect the results more than others, especially conditions that affect a patient’s ability to repeatedly hold his or her breath reproducibly (for example, pulmonary diseases).

Due to the lengthy duration of DENSE breath holds ( 20 seconds) and their limitations  in breath-holding ability,  the breath-hold acquisition protocol was not performed in  patients.   Based on these factors, we  expect that  patients would demonstrate higher variability  in  LV  torsion with  the  breath-hold measures compared to the healthy volunteers  we studied. Hence, the potential reduction in mean variability  when using the respiratory navigator may in fact be higher than the 22% we report from the healthy volunteers in experiment 2.  Nevertheless, the reduction in  LV  torsion variability  patients will  achieve  by  using a respiratory navigator will likely fall between the study’s reported values of 22 and 57%.

3.6  Conclusion

Using a respiratory navigator to  enforce a consistent  end-expiratory position during image acquisition can reduce the variability  in  measured  LV  torsion by

22−57%. Accounting for inconsistent  end-expiratory positions results in favorable

inter-test variability  and reduces required sample  sizes by 80 to 86% compared to previous studies.  Future efforts to measure  LV  torsion should use a respiratory navigator or similar form of consistent respiratory compensation.

CHAPTER 4

OPTIMAL RESPIRATORY NAVIGATOR CONFIGURATION

Adapted from Hamlet SM, Haggerty CM, Suever JD, Wehner GJ, Andres KN, Powell DK, Zhong X, Fornwalt BK. Optimal Configuration of Respiratory

Navigator Gating for the Quantification of Left Ventricular Strain Using Spiral Cine Displacement Encoding with Stimulated  Echoes (DENSE) MRI. Journal of Magnetic Resonance Imaging. 2017. 45(3):796-794

The purpose of this work was  to determine the optimal respiratory navigator gating configuration for the quantification of left ventricular strain using spiral cine displacement encoding with stimulating echoes (DENSE) MRI. In this chapter, we detail  the  different  respiratory  navigator  configurations,   their  advantages and disadvantages,  and the experimental protocol.   This study identifies the optimal respiratory  navigator  configuration  in  adults  and  children  compared to  the

”gold-standard” breath-hold acquisition.

4.1  Synopsis

Purpose: To determine the optimal respiratory navigator gating configuration for the quantification of left ventricular strain using spiral cine displacement encoding with stimulated echoes (DENSE) MRI.

Methods:  Two-dimensional spiral cine DENSE was performed  on a 3 Tesla MRI using two  single-navigator configurations (retrospective, prospective) and a combined dual-navigator configuration in 10 healthy adults and 20 healthy children. The  adults  also underwent   breathhold  DENSE  as   a  reference standard  for comparisons.    Peak left  ventricular  strains,  signal-to-noise ratio  (SNR),  and navigator efficiency were compared. Subjects also underwent dual-navigator gating with and without visual feedback to determine the effect on navigator efficiency.

Results:   There were no differences  in  circumferential, radial,  and longitudinal strains between navigator-gated and breath- hold DENSE (P  = 0.09—0.95)  (as confidence intervals,   retrospective:    [-1.0%–1.1%], [-7.4%–2.0%], [-1.0%1.2%]; prospective:   [-0.6%–2.7%], [-2.8%–8.3%], [-0.3%–2.9%]; dual:    [-1.6%–0.5%], [-8.3%–3.2%], [-0.8%–1.9%], respectively). The dual configuration maintained SNR compared with  breathhold acquisitions (16 versus  18, P = 0.06).   SNR for the prospective   configuration  was   lower  than  for  the  dual  navigator  in   adults (P = 0.004) and children (P < 0.001). Navigator efficiency was higher (P < 0.001) for both retrospective  (54%) and prospective  (56%) configurations  compared with the dual configuration (35%).   Visual feedback  improved the dual configuration navigator efficiency to 55% (P < 0.001).

Conclusion:  When quantifying left ventricular strains using spiral cine DENSE MRI,  a dual navigator configuration results in  the highest SNR in  adults and children.   In  adults, a retrospective  configuration has good navigator efficiency without a substantial drop in SNR. Prospective gating should be avoided because it has the lowest SNR. Visual feedback  represents  an effective  option to  maintain navigator efficiency while using a dual navigator configuration.

Keywords:  Respiratory Navigator, Breath-holds, DENSE, Strain, Signal-to-Noise

Ratio, Navigator Efficiency

4.2  Introduction

Magnetic resonance (MR) can be used to non-invasively assess cardiac function. Displacement encoding with stimulated echoes (DENSE) is an advanced cardiac MR imaging technique that directly measures tissue displacements and can be used to quantify cardiac mechanics, such as myocardial strains and torsion [4, 13].  When combined with clinical risk factors, these measures of cardiac mechanics have been shown to  be  better predictors of mortality  than traditional  measures  of cardiac function, such as ejection  fraction [3].

Compensation for  respiratory  motion  is  an  important  consideration for  all cardiac MR techniques, particularly quantitative imaging sequences like spiral cine DENSE.  DENSE  acquisitions  are  generally  performed  using  end-expiratory

breath-holds  (∼15-20 seconds in duration) [32, 31, 57, 45, 58, 36, 59]; however, this

approach is constrained by the patient’s ability  to breath-hold, which is limited in young subjects and many stages  of advanced  heart disease. Furthermore, short acquisitions  preclude  the   ability   to   capture  more  robust   data,   such  as three-dimensional (3D) DENSE [45, 60, 61], or high resolution imaging [62].

As with  many other cardiac MR sequences,  a respiratory navigator has been used to overcome this time limitation  by allowing the subject to breathe freely but restricting  data  acquisition based on  the  position of  the  diaphragm within  a prescribed ’acceptance’ window [45]. However, unlike some other MR sequences, the navigator echo in the DENSE sequence cannot occur at the beginning of the cardiac cycle, since this would lead to interference with displacement encoding. Instead, the navigator echo must occur at the end of the cardiac cycle, immediately after data acquisition. This creates several options for how the navigator can then be used to either accept or reject the acquired DENSE data (Figure 4.1). For example, a single echo can be used retrospectively or prospectively  to define acceptance of DENSE data from the current  or preceding cardiac cycle, respectively.  Alternatively,  a

dual-navigator configuration can be used, which requires an echo from the current and preceding cardiac cycle to define acceptance of DENSE data (Figure 4.1). Each configuration has distinct advantages and disadvantages.  For example, compared to the single navigator configurations, the dual-navigator configuration has more rigorous criteria for correctly accepting data (Figure  4.2).  However, these strict criteria likely lead to worse navigator efficiency compared to the single navigator configurations (Figure 4.2).

Previous studies using navigator-gated DENSE have reported using a prospective single navigator configuration [45, 61]. However, there has been no formal comparison of the available navigator configurations. Moreover, the accuracy of derived cardiac mechanics and overall image quality for these navigator configurations compared with breath-hold  acquisitions as a reference standard are largely unknown. The purpose of this study was to determine the optimal configuration of respiratory navigator gating for the quantification of left ventricular strain using spiral cine DENSE MRI.

4.3  Methods

4.3.1 Subjects

Ten healthy adults and 20 healthy children with no known history of cardiovascular disease or chronic illnesses and a normal 12-lead electrocardiogram  were prospectively enrolled. The protocol was approved by the local Institutional Review Board and all subjects provided informed consent (or assent/parental consent, as appropriate).

4.3.2 Image Acquisition

Image  acquisition  was   performed on  a  3T  Siemens Tim   Trio   (Siemens Healthcare, Erlangen, Germany) with a 6-element chest coil and a 24-element spine coil. 2D spiral cine DENSE [45, 46] in mid-ventricular short-axis and four-chamber long-axis orientations were separately acquired using breath-holds and retrospective, prospective, and dual navigator gating.  Due to the lengthy breath-hold duration

|  | Navigator Pulse |  |
| --- | --- | --- |
|  | DENSE Imaging |  |
|  | Gates R-R Int. N |  |

Navigator  Gating Configurations

Data are acquired if

R-R

Pro.

Interval N-1

R-R

Interval N

the diaphragm is within the acceptance window:

Pro.

(a) Prospective

DENSE Data Acquisition

Before

Data Acquisition

Retro.

(b) Retrospective

Retro.

DENSE Data Acquisition

After

Data Acquisition

(c) Dual

Retro.

DENSE Data Acquisition

Before and After

Data Acquisition

Pro.

Figure 4.1:  Different navigator gating configurations used to acquire DENSE image data.   The single navigator configurations: (A)  prospective  and (B) retrospective; and the (C)  dual navigator configuration. The red outlines and arrows indicate which navigator pulse(s) is/are gating the DENSE data acquisition.

Prospective

Correctly Discarded

Incorrectly Accepted

Correctly Discarded

Correctly Accepted

Retrospective Incorrectly Accepted

Correctly Discarded

Incorrectly Accepted

Correctly Accepted

Dual

Correctly Discarded

Correctly Discarded

Correctly Discarded

Correctly Accepted

DENSE Data Acquisition

DENSE Data Acquisition

DENSE Data Acquisition

DENSE Data Acquisition

Acceptance Window

Respiratory Signal

Figure   4.2:    Theoretical   example  to   demonstrate   the   disadvantages of  single-  and  dual-navigator   gating  configurations.     The  red  colored text  identifies the incorrectly accepted DENSE data from single-navigator gating configurations (i.e.   the diaphragm is not within  the acceptance  window for the entirety of DENSE data acquisition, but the data are still accepted). The number of Accepted/Discarded data in the example  above illustrate how a dual navigator gating configuration will  discard more data compared  to single navigator gating configurations (retrospective and prospective) and lead to lower navigator efficiency. The red ’x’ or green ’o’ represents the detected diaphragm location being outside or inside the acceptance window, respectively.

(∼20 seconds), breath-hold acquisitions  were not performed in children. The order of the navigator gating configurations was randomized.  Prospective ECG gating was used and the number of cardiac  phases was selected to allow 100-150 ms at the end of the cardiac cycle for heart rate variability.   Acquisition parameters for all scans were: spiral type:  uniform density, interleaves = 6, interleaves per beat = 2, FOV =

360x360 mm2, pixel spacing = 2.8x2.8 mm2, slice thickness = 8 mm, TE = 1.1 ms, TR  = 17 ms, variable flip  angle = 20◦   [46, 63], displacement  encoding = 0.06 cyc/mm   [6],  through-plane  dephasing =  0.08  cyc/mm   [31]  CSPAMM  echo suppression [32], and view sharing.  The temporal resolution was 34 ms, however sliding window view sharing yielded a 17 ms temporal resolution between reconstructed cardiac frames.    Based on  the  DENSE  parameters, acquisition duration for each orientation was 20 heartbeats.

The respiratory navigator was placed over the dome of the liver.  Subjects were

asked to  breathe comfortably  and  a  scout navigator  was  used to   track  the diaphragm.  The navigator acceptance window was  placed so  that  the maximum acceptance window  position  was   located 1-2  millimeters  above   the  subject’s

maximum expiration position.  A navigator acceptance window of ±3 mm (total

range of 7 mm) was used for all navigator gated scans. Navigator efficiency was measured   as  the number of cardiac cycles from which data were acquired and accepted over the total number of cardiac cycles required to complete a scan.

4.3.3 Navigator Feedback

Because  the dual-navigator configuration was  expected  to  decrease  navigator efficiency, we developed and tested a feedback system, which allowed the subject to view their diaphragm position in real-time during image acquisition. The goal was to  compensate  for reduced navigator efficiency in  order to  preserve  the clinical feasibility of DENSE imaging using a dual-navigator configuration.

The feedback system consisted of an angled mirror placed above the patient’s head so that an image of the diaphragm location was viewable on a screen located at the back of the scanner bore. The image of the diaphragm location (respiratory navigator display) was projected from the scanner console’s video feed onto the screen with an MRI-compatible projector. After all other scans were completed,  subjects  used this feedback system, with the dual-navigator gating configuration, to acquire the same short-axis and long-axis images.

This feedback system was also used prior to the breath-hold scan to ensure a consistent  end-expiratory diaphragm location between the  navigator-gated acquisitions and the breath-hold acquisitions. With instruction, the subject exhaled and breath-held in the acceptance window, at which point the navigated scan was halted and the breath-hold acquisition was  immediately performed.  Breath-hold acquisitions were always performed after the navigator-gated  acquisitions that did not involve navigator feedback in order to minimize the potential effect of navigator

feedback on respiratory patterns.

4.3.4 DENSE Post-Processing

All  DENSE images  were analyzed using custom, open-source  MATLAB (The Mathworks Inc, Natick, MA) software, DENSEanalysis [34]. For each set of DENSE images, endocardial  and epicardial boundaries  were drawn on the magnitude image from an end-diastolic and end-systolic frame.  A simplified analysis technique was used to reconstruct the motion field [35]. The displacement-encoded  phase images were unwrapped using a path-following algorithm with  manual selection of seed points.  The resulting Lagrangian displacements underwent  spatial smoothing and temporal fitting as previously  described [36].

Segmental two-dimensional  Lagrangian strains were computed over the cardiac cycle for 6 segments in the short-axis images (radial and circumferential strain) and from the long-axis images (longitudinal  strain).   Cardiac segments  were defined using the American Heart Association 17-segment model. Average peak strains were computed by averaging the strain curves of all the myocardial segments together and  finding  the  peak of  this  average strain  curve.     When  computing  peak longitudinal strain, pixels within 10% of left ventricular longitudinal length from the most basal and apical regions were excluded in order to exclude the increased noise which is typically  observed in the strain curves in those regions. Thickening was defined by convention  as positive strain, whereas shortening  was defined as negative.

4.3.5 Analysis

Mean modified coefficient  of variation (CoV)  [6, 5, 46] was  used to measure agreement  in  strain between different  navigator configurations and breath-holds. The calculation of the CoV is shown below where N is the number of subjects and x1  and x2  are the strain measurements.

Σ   N

C oV =	i=1[St.Dev(x1[i], x2[i])]/N

(4.1)

|Σ

N

i=1

[(x1[i] + x2[i])/2]/N |

Consistent with previous studies reporting CoVs [5, 64, 65, 55, 43], results less than or equal to 20% were considered acceptable.

To compare image quality, signal-to-noise ratio (SNR) was computed  using the DENSE magnitude images at end-systole. SNR was  quantified from the  average myocardial signal and the standard deviation of the noise within an area free from tissue or imaging artifacts [6, 46, 66]. Because the MR signal has a Rician distribution, corrections were applied in order to calculate the true SNR [67].  The measured

standard deviation, σM , was used to compute the true standard deviation, σ, by

2

σ =

4 − π

∗ σM  ≈ 1.526 ∗ σM 	(4.2)

The measured myocardial signal, M , was used to compute the true myocardial signal, S, by

S = √M 2 − σ2	(4.3) SNR was defined as the ratio of the true myocardial signal (S) to the true standard

deviation (σ).

4.3.6 Comparison of Acquisition Configurations

We  compared peak global and segmental  strains (circumferential, radial, and longitudinal) and SNR of the end-systolic DENSE magnitude images between each acquisition technique  (breath-hold and navigator gating) in adults.  Bland-Altman analyses  [38], CoV [5], and 95% confidence  intervals (CI)  were used to measure agreement in strain between the separate navigator configurations and breath-holds. A   paired  Student’s   t-test   was   used to   compare strains  between navigator

configurations and breath-holds. We  also compared SNR and navigator efficiency between all navigator configurations (dual, retrospective, and prospective) in adults and children using a one-way repeated  measures ANOVA  with  post-hoc analyses

and Bonferroni correction.    All  data  are presented as  mean ± one standard

deviation. Significance  was defined as p < 0.05.

4.4  Results

Ten healthy adults (Age: 23 ± 3 years, 40% female) and 20 healthy children (Age:

13 ± 3 years, 45% female) were enrolled in the study. DENSE data were successfully acquired in all subjects except for one child who could not complete the study protocol

due to an erratic respiratory pattern.

4.4.1 Average Peak Strains

Average peak left ventricular strains are shown in Table 4.1.  There were no significant   mean differences in  circumferential,  radial,  and  longitudinal  strain between the  dual,  retrospective, and  prospective  navigator  configurations and breath-holds in  adults (Figure  4.3).    Compared to  breath-holds, all  navigator configurations had  a  CoV  of  less   than  20% for  circumferential,  radial,  and longitudinal strain in adults (Figure 4.3).  The differences  in strain are listed as confidence intervals in Table 4.2. Peak segmental left ventricular strains are shown in Table B.1 in Appendix A).  There were no significant  differences in segmental strain between the navigator configurations and breath-holds except for radial strain from  the  prospective   configuration  (p  =  0.002, Table  B.2).     Compared to breath-holds, all navigator configurations had CoVs of less  than 20% except for radial segmental strain (19-28%) (Table B.2).

Table 4.1: Average strains for different acquisition techniques.

Circumferential  Strain (%)

Mean ± Std.  Dev.

Adults 	Children

Breath-hold	-17 ± 2		– Retrospective	-18 ± 2	-19 ± 2

Prospective	-17 ± 3	-18 ± 2

Dual	-18 ± 3	-20 ± 2

Radial Strain (%)

| Breath-hold Retrospective Prospective Dual | 30 ± 10 26 ± 9 31 ± 7 27 ± 9 | – 30 ± 9 26 ± 12 27 ± 12 |
| --- | --- | --- |
| Longitudinal Strain (%) |  |  |
| Breath-hold Retrospective Prospective Dual | -14 ± 2 -14 ± 2 -13 ± 2 -14 ± 2 | – -14 ± 2 -14 ± 2 -14 ± 2 |

Table 4.2: CI  Results for Differences in Strain Between Navigator  Gating and Breathhold DENSE.

Circumferential  Strain (%)

95% LoA 	p-value

Retrospective–Breath-hold	[-1.0–1.1]	0.95

Prospective–Breath-hold	[-0.6–2.7]	0.17

Dual–Breath-hold	[-1.6-0.5]	0.29

Radial Strain (%)

| Retrospective–Breath-hold | [-7.4–2.0] | 0.23 |
| --- | --- | --- |
| Prospective–Breath-hold | [-2.8–8.3] | 0.29 |
| Dual–Breath-hold | [-8.3-3.2] | 0.34 |
| Longitudinal Strain (%) |  |  |
| Retrospective–Breath-hold | [-1.0-2.0] | 0.85 |
| Prospective–Breath-hold | [-0.3-2.9] | 0.09 |
| Dual–Breath-hold | [-0.8–1.9] | 0.38 |

Retro vs Breath-hold	Pro vs Breath-hold

Dual vs Breath-hold

Ecc Retro − BH (%)

Ecc

Mean CoV = 5%

6

4

2

0

−2

−4

p = 0.95

−6

2.9

0.0

−2.9

Mean CoV = 8%

5.6 1.1

−3.4

p = 0.17

6

Ecc Pro − BH (%)

4

2

0

−2

−4

−6

Mean CoV = 6%

2.4

−0.5

−3.4

p = 0.29

6

Ecc Dual − BH (%)

4

2

0

−2

−4

−6

−22  −20  −18  −16  −14  −12

Mean (%)

−22 −20 −18 −16 −14 −12

Mean (%)

−22  −20  −18  −16  −14

Mean (%)

Mean CoV = 15%

p = 0.23 10.3

−2.7

−15.7

20

Err Retro − BH (%)

10

Mean CoV = 14%

p = 0.29 18.0 2.8

−12.5

20

10

Mean CoV = 17%

p = 0.34 13.3

−2.6

−18.4

20

10

Err Pro − BH (%)

Err Dual − BH (%)

Err 	0 	0 	0

−10

−10

−10

−20

10 	20 	30 	40 	50

Mean (%)

−20

10 	20 	30 	40 	50

Mean (%)

−20

10 	20 	30 	40 	50

Mean (%)

3.2

0.1

−3.0

p = 0.85

6

Ell Retro − BH (%)

4

2

Ell

0

−2

−4

−6

Mean CoV = 7%

Mean CoV = 11%

6

Ell Pro − BH (%)

4

2

0

−2

−4

p = 0.09

−6

6

5.7

Ell Dual − BH (%)

4

1.3 	2

0

−3.0 	−2

−4

−6

Mean CoV = 8%

−18 	−16 	−14 	−12 	−10

Mean (%)

−16 	−14 	−12 	−10

Mean (%)

−16 	−14 	−12

4.2

0.5

−3.1

p = 0.38

Mean (%)

Figure  4.3:   Bland-Altman plots of average peak circumferential  (Ecc), radial (Err), and longitudinal (Ell)  strains for retrospective, prospective, and dual navigator gating vs breath-hold.

Table 	4.3: 	Signal-to-noise	ratios	for 	different	navigator 	gating configurations in adults and children.

Adults 	p-value 	Children 	p-value

Breath-hold	18 ± 8

Retrospective	15 ± 6*†

< 0.001

–

22 ± 6‡

< 0.001

Prospective	13 ± 5* 	20 ± 8‡

Dual	16 ± 7	27 ± 9

* p < 0.05 vs. breath-hold; †  p < 0.05 vs. prospective; ‡  p < 0.05 vs. dual

4.4.2 Signal-to-Noise Ratio

In  adults,  single navigator configurations had a  17-28% reduction in  SNR compared to breath-hold DENSE (Table 4.3).   There was  no difference  in SNR between the dual navigator configuration and breath-hold DENSE (p  = 0.06). Among navigator configurations, dual and retrospective  navigator configurations were comparable and both had better SNR (23% and 15%, respectively) compared to the prospective configuration (p = 0.02, p = 0.004, respectively).

In children,  SNR also differed based on navigator gating configuration (Table 4.3). The dual navigator configuration had the highest SNR compared to the retrospective (23% higher, p < 0.001) and prospective  (35% higher, p < 0.001) configurations. There was no difference in SNR between the retrospective and prospective navigator configurations (p = 0.15).

4.4.3 Navigator Efficiency

For adults and children combined, there were significant differences in navigator efficiency between navigator configurations (p  < 0.001, Table 4.4).     The retrospective  and prospective  navigator configurations had higher navigator efficiencies than the dual navigator configuration by an average of 54% (p < 0.001) and 60% (p < 0.001), respectively. Using visual feedback with the dual navigator configuration improved navigator efficiency by 57% (p < 0.001) compared to the

Table 	4.4: 	Navigator 	efficiencies	for 	different	navigator 	gating configurations in adults and children.

Pooled	p-value 	Adults 	Children

Retrospective	54 ± 15* Prospective	56 ± 15*

< 0.001

48 ± 15*	57 ± 16*

52 ± 17*	58 ± 14*

Dual	35 ± 13	31 ± 16	37 ± 11

Dual Feedback	55 ± 16*	67 ± 11*	48 ± 15*

* p < 0.05 vs. dual (without  feedback)

dual configuration without  feedback  and resulted in comparable efficiency to the single navigator configurations. The scan times mirrored the navigator efficiency results.  For example, scan times for adults were, on average,  64, 37, 37, and 25 seconds for  the  dual,  retrospective, prospective   configurations, and  the  dual navigator configuration with feedback, respectively.

4.5  Discussion

The use of respiratory navigated acquisitions for spiral cine DENSE extends the potential utility  of the technique by removing restrictions on patient breath-holding abilities and allowing for high resolution [62] and/or three dimensional [45, 60] data collection. While previous  studies have used a respiratory navigator with DENSE, the optimal configuration for gating and how it might differ among subjects has been largely unexplored. This study addressed these knowledge  gaps by showing that:  1) left ventricular peak strains were not different between breath-held and navigator- gated DENSE acquisitions; 2) SNR was reduced with single navigator configurations, but not the dual configuration, compared to breath-held acquisitions; 3) the SNR benefit of the dual navigator configuration  was offset by reduced navigator efficiency compared to single navigator configurations, but visual navigator  feedback maintained clinically acceptable  efficiencies for the dual navigator acquisition.  The following paragraphs explore each of these findings in greater detail.

There were no significant  mean differences  and good paired agreement  of all

peak strains between retrospective,  prospective,  and dual navigator configurations and breath-holds in adults. This finding agrees with the prior work by Zhong et al., which compared segmental strains from navigator-gated 3D DENSE to breath-hold

2D DENSE [45] and similarly reported clinically acceptable agreement. Our study extends this  work by demonstrating that  the agreement exists not  only for prospective  navigator-gating used by Zhong et al., but for retrospective  and dual navigator-gating configurations as  well.   Demonstrating this agreement  of strain valuesa primary endpoint for most DENSE acquisitionshas pragmatic value by ensuring that data from acquisitions with differing respiratory compensation can be readily compared.

Our results demonstrate that using a single navigator configuration resulted in significantly lower  SNR compared to breath-hold DENSE acquisitions. While this result was only demonstrated in adults because of the prohibitively long breath-hold duration in children, it is reasonable to assume that a similar trend holds in children  as well. Among the navigator configurations, the dual configurations provided the best SNR as it was superior to prospective navigator gating in both adults and children, had better SNR than retrospective  gating in children, and resulted in comparable SNR to breath-hold DENSE.

Differences in SNR among the different  acquisitions are likely attributable  to heart rate and respiratory variability.  The breath-hold acquisitions had the shortest acquisition time,  with  presumably less  physiologic variability.     Also,  the  dual navigator configuration had the most stringent  acceptance  criteria,  which likely minimized the effects of respiratory variability  during acquisition compared to the other configurations. This reasoning is supported by previous studies, which have reported associations between consistent diaphragm position during navigator-gated acquisitions and  improved SNR  [68, 69].    In  both  adults  and  children,  the prospective   navigator  configuration  had  the   lowest  SNR  of   all   navigator

configurations.   The  observed difference in  SNR between the  single navigator configurations was  perhaps unexpected  given the theoretical similarities in  their design and function.   However, these differences  are similarly attributable  to the effects of variability:   for the retrospective  navigator, the interval between the R wave and the navigator echo is fixed, whereas heart rate changes during the scan will  affect the interval between the navigator echo and the succeeding R wave in prospective  gating, increasing the likelihood of respiratory variability  during that interval.    Based on these findings, the use of prospective  navigator gating for DENSE should be avoided.

Notably, a previous study compared SNR of a 2D steady-state free precession sequence between dual navigator-gated and breath-hold acquisitions and found that end-systolic myocardial SNR for breath-hold acquisitions  was 23% lower than the dual navigator configuration in adults [70]. This finding contrasts with our data in which the dual navigator configuration was statistically comparable to breath-hold. However, the previous study had substantially different imaging parameters between their dual navigator-gated acquisition and the breath-hold acquisition, which likely accounted for the observed SNR differences [70].

Although the purpose of this study was to determine the optimal navigator gating strategy, it is worth noting that SNR was higher for children than it was for adults. The difference in SNR between adults and children is likely related to the smaller body habitus of children, which results in a shorter distance between the MRI coils and the heart. Moreover, adults likely have more adipose tissue, which could also lead to lower SNR. Ultimately, these SNR differences may lead to differences in inter-test reproducibility between adults and children.

As expected, single navigator gating configurations resulted in better navigator efficiency compared to dual navigator gating, due to the additional acceptance criteria constraints of the dual navigator.  Simply put, more data are discarded with dual

navigator gating, leading to prolonged  scan time.  Previous studies using a single- navigator configuration with the same size acceptance window (±3 mm) reported navigator gating efficiencies ranging from 20 to 48% [45, 68, 71]. Compared to these

studies, we observed slightly better single-navigator  efficiencies of 48 to 52% in adults and 57 to 58% in children. The dual navigator efficiency was comparable  to results from previous studies [70].

To potentially offset this reduced efficiency, we evaluated the effect of providing the subject with visual feedback of the diaphragm position, which has been shown to considerably improve navigator efficiency compared to traditional free-breathing acquisitions [68]. We found that using visual feedback during dual navigator gated acquisitions improved navigator efficiency compared to the dual configuration without feedback and resulted in comparable efficiency to the single navigator configurations. The improvement with feedback was not uniform across adults and children (i.e., the improvement in kids was not as substantial), perhaps reflecting the superior ability of the adults to hold their breath within the acceptance window. Alternatively, the difference may be indicative of the non-intuitive nature of the respiratory navigator display and the differential abilities of adults and kids to quickly learn and use it. Efforts are ongoing to instead transform this image to a more kid-friendly video game design to improve usability [30].

The  increased scan time  associated  with  the  dual  navigator  configuration presents an obvious trade-off with improved SNR for its utility  in a clinical setting where time is a critical consideration. In adults, given the  minimal difference in SNR between the dual and retrospective  navigator configurations, the substantial drop in efficiency with the dual navigator may not be justified. In children, however, the SNR benefit with  dual navigator gating is more substantial and warrants consideration to ensure high quality data. Hence the demonstrated improvement in navigator efficiency by providing the subject with visual feedback is an important

finding because  it provides one option for compromise: achieving improved SNR while  approximately maintaining  scan time  compared to  other  navigator configurations.

This study was performed using 2D DENSE. However, given the similarity in the fundamental sequence designs  of 2D and 3D DENSE, the results are applicable to

3D DENSE acquisitions   as well.  In fact, given the longer time generally required for 3D data acquisition, respiratory compensation/navigation is essential,  so these results are highly relevant. Specifically, our navigator efficiency findings agree with reported efficiencies from a previous study using 3D DENSE [45]. Also, while absolute magnitudes of SNR may differ with more data acquired, there is no reason to suspect that the relative differences in SNR would change between different navigator gating strategies when applying these results to 3D DENSE.

A limitation  of this study is the potentially limited power for detecting small strain  differences between navigator  configurations and  breath-hold  DENSE. However, our study had 80% power to detect a difference  of 1.5% between the retrospective  navigator configuration and breath-holds.   This  1.5% difference is smaller than the typical inter-test limits of agreement of circumferential strain [6]. Moreover,  even if the strains from the  prospective  navigator configuration, which had the worst agreement,  are in fact different from the breath-hold technique, the conclusions of  the  study  would  not  change as   the  prospective  navigator configuration  was separately found to be sub-optimal based on SNR.

Another limitation  was the lack of breath-hold data for the  pediatric subjects. The DENSE acquisition required 20 heartbeats. The required breath-hold time was further extended by the use of a navigator-gated  pre-scan to ensure that breath-holds were performed  at the same diaphragm position as the navigator-gated  scans. This duration was viewed to be prohibitively long for pediatric subjects, and so no breath- hold DENSE data were acquired in these cases. The equivalence of DENSE-derived

strains between breath-hold and navigator sequences was demonstrated   in adults. Since children did not undergo breath-hold DENSE, we must caution future studies regarding these strain results as they apply to children. However, since the relative SNR and navigator efficiency results from navigator gating were similar to those in adults, we would expect potential differences to be small. Furthermore, the primary objective of the study was to identify optimal navigator configurations for DENSE, so the lack of breath-hold data in children is a minor limitation.

A third  limitation  of this study is the lack of assessment  of clinical patients. Cardiac patients, who routinely undergo MR imaging and who may have limited ability to hold their breath, may not be able to perform the lengthy breath-hold scan and may not achieve  as high navigator efficiency when performing a dual navigator scan with feedback. However, since this population is more likely to undergo DENSE MR imaging than this study’s volunteer subjects, it would be beneficial to determine whether the results remain the same. For example, it may be important to use dual navigator gating, even at the expense of navigator efficiency, to achieve higher SNR, since SNR is commonly lower in the clinical patient population compared to healthy volunteers.

In conclusion, for spiral cine DENSE acquisitions, respiratory navigator gating and breath-hold acquisitions yield comparable values of left ventricular peak strains. However, differences in signal-to-noise ratios and navigator efficiencies were observed among the different navigator gating configurations, which warrant consideration in clinical  and  research protocol  design.    In  adult  subjects, the  dual  navigator configuration produced the best SNR, although only slightly better than the single retrospective navigator, which produced acceptable SNR and therefore may be used to maintain good efficiency.   For children, the benefit of a dual navigator configuration for improved SNR was  even  more apparent, but  resulted in a considerable  drop in  scan efficiency.  The prospective  navigator resulted in  the

poorest SNR and  should be  avoided.    The  use of  visual  navigator  feedback represents an effective option to maintain navigator efficiency while using the dual navigator in children (and adults).

CHAPTER 5

INTERACTIVE  FEEDBACK GAME DESIGNED TO  IMPROVE NAVIGATOR EFFICIENCY

Adapted from Hamlet SM, Haggerty CM, Suever JD, Wehner GJ, Grabau JD, Andres KN, Vandsburger MH, Powell DK, Sorrell VL, Fornwalt BK. An interactive videogame designed to improve respiratory navigator efficiency in children undergoing cardiovascular magnetic resonance.  Journal of Magnetic Resonance Imaging. 2016. 18:54

5.1  Synopsis

Background:   Advanced cardiovascular magnetic resonance  (CMR)  acquisitions often require long scan durations that necessitate respiratory navigator gating. The tradeoff of  navigator  gating  is  reduced scan efficiency, particularly  when the patient’s breathing patterns are inconsistent,  as is commonly seen in children. We hypothesized that   engaging pediatric  participants  with   a  navigator-controlled videogame  to help control breathing patterns would improve  navigator efficiency and maintain image quality.

Methods:  We  developed custom software that processed the Siemens respiratory navigator image in real-time during CMR and represented diaphragm position using a cartoon avatar, which was projected to the participant  in the scanner  as visual feedback.  The game incentivized children to  breathe such that  the avatar was

positioned within  the navigator acceptance  window (±3 mm) throughout image

acquisition.  Using a 3T Siemens Tim Trio,  50 children (Age:  14 ± 3 years, 48%

female)  with   no  significant   past  medical  history   underwent   a  respiratory navigator-gated 2D  spiral  cine displacement   encoding with  stimulated   echoes (DENSE) CMR acquisition first with no feedback (NF) and then with the feedback

game (FG).  Thirty   of  the  50 children were randomized to  undergo extensive off-scanner training with the FG using a MRI simulator, or no off-scanner training. Navigator efficiency, signal-to-noise ratio (SNR), and global left-ventricular strains were determined for each participant and compared.

Results: Using the FG improved average navigator efficiency from 33 ± 15 to 58 ±

13% (p < 0.001) and improved SNR by 5% (p = 0.01) compared to acquisitions with NF. There was no difference in navigator efficiency (p = 0.90) or SNR (p = 0.77) between untrained and trained participants for FG acquisitions. Circumferential and radial strains derived from FG acquisitions  were slightly reduced compared to NF

acquisitions (16 ± 2% vs 17 ± 2%, p < 0.001; 40 ± 10% vs 44 ± 11%, p = 0.005,

respectively). There were no differences in longitudinal strain (p = 0.38). Conclusion: Use of a respiratory navigator feedback game during navigator-gated CMR improved navigator efficiency in children from 33 to 58%.	This improved efficiency was  associated  with  a  5% increase in  SNR for  spiral  cine DENSE. Extensive  off-scanner  training  was  not  required to  achieve  the improvement  in navigator efficiency.

Keywords:  Pediatrics, Respiratory navigator, Navigator efficiency, Image quality, Cardiovascular magnetic resonance

5.2  Background

Cardiovascular  magnetic  resonance (CMR) can be used to non-invasively assess heart function. In the clinical setting, CMR techniques play an important role in the diagnosis and monitoring of the complex anatomy and physiology of congenital and acquired heart diseases. Moreover, there is a considerable body of pre-clinical research devoted to the development  and evaluation of new, advanced imaging techniques, such as 3D displacement encoding with stimulated echoes (DENSE) [45], 3D steady state free precession [72], and 4D flow imaging [73].  These new techniques have demonstrated ability  in distinguishing normal and pathological tissue deformation and blood flow and may become beneficial tools in the diagnosis and management of heart disease. Many of these clinical and pre-clinical techniques require scan durations that exceed patients’ ability to hold their breath.

End-expiratory breath-holds are used by  many CMR  sequences  in  order to minimize respiratory-motion artifacts.   However, requiring subjects to hold their breath introduces significant limitations on the duration of data acquisition or the quality  of the acquired images, particularly  for young children or patients  with advanced  disease. A  common alternative  is respiratory navigator gating, which works by  measuring  the diaphragm position during normal breathing and only acquiring data when the diaphragm is within  a pre-defined acceptance  window (Figure 5.1a).   The trade-off of navigator gating is significantly increased scan duration because of poor navigator efficiency. For example, previous CMR studies have  have  reported respiratory  navigator  efficiencies  of  20 to  45% in  adults [71, 68, 69, 74]. This poor navigator efficiency lengthens the duration of currently used clinical imaging and limits  clinical feasibility of emerging  advanced imaging techniques.

Navigator efficiency is typically poor because breathing patterns can be erratic

[21, 22, 23] and the patient is generally unaware of the desired acceptance window

location.   Providing the patient  with  visual feedback  of the diaphragm position during  CMR  (”navigator  feedback”)has been shown to  improve  breathing consistency and scan efficiency  in adults [68, 21]. For example, studies have shown efficiency improvements  up to 29% (absolute) compared to traditional acquisitions without  feedback [68, 69].  Importantly,  these previous studies have demonstrated that  image  quality from navigator feedback  acquisitions is similar to acquisitions without feedback [68, 69]. The potential to achieve similar benefits using navigator feedback with pediatric participants has not been explored. Given the challenge of keeping these participants  still  and  motionless for  long periods of  time,  this improved efficiency could have substantial clinical benefit.

Most previous studies involving navigator feedback simply utilized the built-in navigator display. One previous study evaluated a custom videogame interface in a study  of  adults  for  increasing navigator  efficiency [69].     Such   an  interface theoretically combines the benefits of visual feedback with an intuitive and engaging design for the user–attributes that are highly desirable for pediatric scanning. Thus, the present study sought to extend and tailor this paradigm specifically for children by  providing  navigator  feedback in  the  form  of  an  interactive,  kid-friendly videogame.  Moreover, this  study sought  to  test this  design using DENSE, an imaging technique that can be used to quantify advanced measures of function such as  cardiac strains.  We  hypothesized that  navigator feedback using an interactive videogame  during CMR  would improve  navigator efficiency and maintain image quality and strains in children.

5.3  Methods

5.3.1 Feedback videogame

A  navigator feedback  videogame  (FG),  called ”Bubble  Gulp”,  was  developed using MATLAB (The Mathworks Inc, Natick, MA).  Each frame of the navigator

Cropped Siemens

Syngo Navigator Image

Exhale Direction

Acceptance Window

Inhale Direction

Diaphragm

A	Position	B

Example Frame of

Feedback Videogame Interface

Bubble Gulp

Instructions

- How to move fish

within acceptance window

Figure  5.1:  Feedback videogame.  (A)  Cropped version of the Siemens Syngo navigator image that  was  processed  in real-time during CMR acquisition to yield the feedback videogame. (B)  Example frame of the navigator feedback videogame interface, which was shown to the child during CMR (yellow overlay text was not shown to the child).

image provided within  the  Siemens Syngo user-interface (Siemens Healthcare, Erlangen, Germany) (Figure 5.1a) was  captured using an Epiphan DVI2USB 3.0 (Epiphan Systems Inc.,  Palo Alto,  California)  frame grabber and processed  in real-time during  CMR  to  yield  a kid-friendly  representation of  the  diaphragm position (Figure 5.1b).    Navigator image processing  was  performed using an externally connected laptop running Windows 7 with an Intel Core i7 processor and

16 GB of RAM.  The FG interface  was  then projected to the participant  in the scanner  using an angled mirror  and a magnetic resonance  compatible projector (Figure 5.2).

The diaphragm position relative  to the acceptance  window (Figure 5.1a) was represented by the vertical position of a fish character relative to parallel green lines containing scrolling dots, representing bubbles (Figure 5.1b). The objective of the game was  to  control the fish’s vertical position, which was  updated with  each navigator pulse,  so it would ”gulp”  bubbles and acquire points. To incentivize slow, stable breathing, point values increased as the fish spent more time within the green

MRI Scanner

Mirror

MRI Scanner

Mirror

Feedback

Videogame

Figure  5.2:  MRI Feedback Setup Feedback  videogame  was shown to children during CMR with an angled mirror and MR-compatible projector.

lines, instead of frequent short-duration breath-holds. However, prior to any use of the FG, children were instructed to not hold their  breath for an uncomfortable amount  of time and to breathe when needed. Finally, the FG interface displayed text to instruct children how to adjust their breathing in order to place the fish in between the green lines (Figure 5.1b).

5.3.2 Participants

Fifty children with no significant past medical history were recruited to participate in the study. Participants were recruited from the broader clinical community based out of our university medical center using a wide range of participant recruitment services provided by the University of Kentucky Center for Clinical and Translational Science. All participants were screened with a 12-lead ECG prior to imaging to rule out arrhythmias. The local Institutional Review Board at the University of Kentucky approved the study protocol and all participants and legal guardians provided written informed consent or assent.

5.3.3 Imaging

All  imaging was performed using a 3T Siemens Tim Trio (Siemens Healthcare, Erlangen, Germany) with a 6-element  chest coil and a 24-element  spine coil.  For

each participant, navigator-gated 2D spiral cine DENSE CMR [45, 46] images from mid-ventricular, 4-chamber, basal, and apical image orientations were separately acquired with  no feedback  (NF)  and then while using the FG. No instructions regarding breathing were given for  the NF  acquisitions, thus participants were allowed to breathe naturally. Between acquisitions with NF and those with the FG, each participant underwent two 30-heartbeat  practice  scans to familiarize himself or herself with the FG.

DENSE  imaging parameters included:   number of  spiral  interleaves  = 12, interleaves per beat = 2, FOV=  360  360 mm2, pixel spacing = 2.8  2.8 mm2,slice thickness = 8 mm,  TE  = 1.4 ms, TR  = 17 ms, variable flip  angle = 20◦, displacement encoding = 0.06 cyc/mm [6], through-plane  dephasing = 0.08 cyc/mm [31], CSPAMM echo suppression  [32], view sharing and a dual- navigator strategy

[30] with  an acceptance  window size of ±3 mm.    For  each  cardiac cycle, the

navigator echo  occurred immediately after data acquisition.  The dual-navigator strategy required the diaphragm position to be within  the  acceptance window for both the preceding and current  cardiac cycles  in order for data to be accepted. Prospective  ECG gating was  performed and 11–25  cardiac phases were  acquired depending on participant  heart rate.  As a result of the  imaging parameters, each complete image acquisition required 38 heartbeats that  satisfied the  navigator gating criteria.

Due to erratic respiratory patterns or participant  movement, image acquisition can be difficult to complete in children in a reasonable amount of time with NF. As scan session duration increases, the likelihood of patient movement also increases, so we defined criteria for maintaining a target scan protocol duration of 30 min.  We defined image acquisition  as incomplete (data not acquired) following 192 heartbeats without a completed image acquisition. Progressing  past 192 heartbeats for a 38- heartbeat scan is equivalent to achieving  less than 20% navigator efficiency, which

is worse than previously reported NF values [71, 68, 69, 74]. Once any NF image acquisition was marked as incomplete, we proceeded to the FG acquisitions. If a participant moved, the number of acquired image orientations  was reduced from four (mid, 4ch, base, apex) to two (mid, 4ch) to ensure at least two images were acquired with NF and FG.

5.3.4 Calculation of cardiac strains from DENSE

DENSE images were analyzed using DENSEanalysis [34], a custom, open-source MATLAB (the  Mathworks Inc,  Natick,  MA)  software.    To  delineate the myocardium, endocardial and epicardial boundaries  were manually drawn on the DENSE magnitude  image using an end-systolic and end-diastolic frame.   The motion field was reconstructed   using a simplified analysis technique [35].  Using manual  selection of  seed   points,  which  indicated  unwrapped phase data,  a path-following algorithm was used to unwrap the displacement-encoded  phase data. Temporal fitting  and spatial smoothing was  applied to the resulting Lagrangian displacements  as previously  described [36].

Two-dimensional segmental Lagrangian strains  were quantified  from  the smoothed trajectories over the entire cardiac cycle.   Radial and circumferential strain  was  computed for  6  myocardial segments  of  the  short-axis images and longitudinal strain was computed from the long-axis images. The strain curves of all the cardiac segments were averaged into a single average curve. Global peak strain was  quantified by  averaging the  strain  curves  from  each  slice and finding  the resulting peak strain of this curve. When computing peak longitudinal strain, pixels within  10% of left  ventricular longitudinal  length of the most basal and apical regions were excluded due to increased noise typically observed in the strain curves in those regions. Peak strain was defined  as a positive for thickening (radial strain) and negative for shortening (circumferential and longitudinal strain).

5.3.5 Analysis

This study measured navigator efficiency and heart rate during image acquisition and used image signal-to-noise ratio (SNR) of the end-systolic DENSE magnitude image as a measure of image quality. Navigator efficiency  was defined as the ratio of the number of heartbeats for which image data were accepted to the total number of heart beats required to complete the image acquisition. To compare image quality, signal to noise ratio (SNR) was calculated  for each cardiac phase of each DENSE magnitude image. SNR was computed  from the average myocardium  signal and the standard deviation of the signal (noise) within an area with no signal (free from tissue and imaging artifacts). Due to the Rician distribution of the MR signal, corrections were applied to the measured standard deviation (σM  in Equation 5.1) and measured myocardial signal (M in Equation 5.2) to compute the true SNR [46, 6, 67]. The SNR was defined as the ratio of the true myocardial signal to the true standard deviation.

2

σ =

4 − π

∗ σM  ≈ 1.526 ∗ σM 	(5.1)

S = √M 2 − σ2	(5.2)

For incomplete NF image acquisitions (satisfied stoppage criterion),  navigator efficiency and heart rate measurements were computed based on the partial data that were acquired.

5.3.6 Training

Off-scanner training has been used by other investigators to ensure participants are comfortable and understand a navigator feedback interface before entering the magnet [68]. We wanted to determine the efficacy of off- scanner training with the FG on navigator efficiency, image quality, and heart rate. Thus, 30 of the 50 enrolled

participants were randomized into equal groups to either receive extensive off-scanner training or no off-scanner training prior to scanning; thus, the groups were referred to as ’trained’ and ’untrained.’  As mentioned above, all subjects (including trained and untrained participants) underwent minimal training in the scanner, which was defined as two 30-heartbeat  practice  scans prior to FG acquisitions. The remaining 20 participants also received off-scanner training, but they were not included within the trained subgroup for analysis because they were not randomized to this treatment.

Each trained participant was introduced to the FG using an MRI simulator prior to the formal study. The MRI simulator utilized a PrimeSense Carmine 1.09 (Prime- Sense, Tel Aviv, Israel) 3D camera to precisely measure the chest wall and abdomen excursion  as a proxy for diaphragm translation [75, 76].  Each participant  had to complete goal-based training before advancing to CMR scanning. Training time was recorded for all trained participants. The training protocol is described in detail in Appendix B.

5.3.7 Statistics

Statistical analyses  were completed using R version 3.2.2 (R  Foundation for Statistical Computing, Vienna, Austria).     All  continuous measurements  were reported as mean ± standard deviation. Navigator efficiency, SNR, heart rate, and global left ventricular strains were tested for normality using a Shapiro-Wilk test. Average navigator efficiency, SNR, heart rate, and strain were compared between NF and FG acquisitions using a paired student’s t-test or Wilcoxon Signed-Rank test when appropriate, and compared between untrained and trained groups using a student’s t-test or Mann-Whitney U test when appropriate. To determine whether age had  an  effect on  navigator  efficiency, age was  correlated with  navigator efficiency for both NF and FG acquisitions.

5.4  Results

Fifty-six children were prospectively  enrolled. Six children were excluded from the study due to either being uncomfortable in an MRI scanner, having premature ventricular  contractions, having ECG-monitoring equipment  fail,  or consistently moving during scanning. Thus, this study reported data on 50 children (Age: 14 ±

3 years,  48% female) with  no significant  past medical history, which included a subset of 30 children randomized to either the off-scanner trained (n = 15; Age: 15

± 3 years, 47% female) or untrained (n = 15; Age: 13 ± 3, 66% female) groups. All

trained  participants  successfully  completed off-scanner training  and  the  mean training duration was 11 ± 2 min.  The prescribed  stoppage criterion for the NF scans was  met in  11 cases,  resulting in  fewer completed NF  images for  those

participants. Additionally, four participants moved during scanning, which included two during NF scans and two during FG scans, resulting in the completion of the abridged imaging protocol, as described  in the methods.

5.4.1 Navigator Efficiency

Using the FG significantly improved average navigator efficiency compared to NF (58 ± 13% vs 33 ± 15%, p < 0.001, Figure 5.3a).  Average navigator efficiency was not correlated with age for either NF or FG image acquisitions (r = 0.07, p =

0.63; r=0.14, p = 0.32, Figure 5.3b). There was no significant difference in average navigator efficiency between untrained and off-scanner trained groups for FG image

acquisitions (57 ± 17% vs 57 ± 11%, p = 0.90, Figure 5.4).

5.4.2 SNR

Use of the FG significantly improved SNR compared to NF (22 ± 6 vs 21 ± 6, p

= 0.01, Figure 5.5). There was no significant difference in SNR between untrained and off-scanner trained groups for FG images (22 ± 6 vs 21 ± 6, p = 0.77).

Average Navigator Efficiency (%)

100

80

p < 0.001

100

80

NF, r = -0.07, p = 0.63

FG, r = 0.14, p = 0.32

Average Navigator Efficiency (%)

60 	60

40 	40

20 	20

0  A 	0  B

No

Feedback

Feedback

Game

8 	10 	12 	14 	16 	18

Age (years)

Figure 5.3: (A)  Average navigator efficiency for No Feedback and Feedback Game  image  acquisitions.   Use of the feedback game significantly increased navigator efficiency compared to no feedback.  The solid red line indicates the mean of each group.  (B)  Average  navigator  efficiency vs age for No  Feedback (NF) and Feedback Game (FG) image acquisitions. There was no correlation between navigator efficiency and age for either no feedback (r = -0.07, p = 0.63) or feedback game (r = 0.14, p = 0.32) acquisitions. The solid lines indicate the line of best fit for each group.

100

Feedback Game

p = 0.90

Average Navigator Efficiency (%)

80

60

40

20

0

Untrained	Trained

Figure 5.4:  Average navigator efficiency for Off-scanner Trained and Untrained   groups.   There was  no significant  difference  in  navigator efficiency between untrained and off-scanner trained groups for feedback game acquisitions. The solid red line indicates the mean of each group.

45

p = 0.01

40

35

30

SNR

25

20

15

10

5

0	No

Feedback

Feedback

Game

Figure 5.5:  SNR  for all No Feedback and Feedback Game images. Use of the feedback game resulted in significantly increased SNR compared to no feedback.

110

Mean Heart Rate (bpm)

100

90

80

25

p < 0.001	p = 0.30

Heart Rate Variability (bpm)

(Standard Deviation)

20

15

70	10

60

5

50

40

No

Feedback

Feedback

Game

0

No

Feedback

Feedback

Game

Figure 5.6: Mean and standard deviation of heart rate for No Feedback and Feedback Game acquisitions. Use of the feedback game resulted in significantly higher heart rate compared to no feedback. There was no significant difference in standard deviation of heart rate between no feedback and feedback game acquisitions. The solid red line indicates the mean of each group.

5.4.3 Heart rate

On average, heart rate during FG scans was slightly higher than NF acquisitions (75 ± 13 vs 72 ± 12 bpm, p < 0.001, Figure 5.6), but there were no differences in the standard deviation of heart rate (5.9 ± 2.2 vs 6.1 ± 3.9 bpm, p =0.30, Figure 5.6b). Heart rate was similarly elevated during FG acquisitions in both the untrained and off-scanner trained groups compared to NF acquisitions (p < 0.001 and p = 0.03, respectively, Table 5.1).

5.4.4 Strain

Global circumferential and radial  strains derived from FG  acquisitions  were slightly lower in magnitude compared to NF acquisitions (16 ± 2% vs 17 ± 2%, p <

0.001; 40 ± 10% vs 44 ± 11%, p = 0.005, respectively,  Table 5.2). There were no

77

Table 5.1: Average Heart  Rate for Off-scanner Trained and Untrained  groups.

HeartRate (bpm) 	Trained	Untrained

|  | No Feedback | Feedback Game | p-value | No Feedback | Feedback Game | p-value |
| --- | --- | --- | --- | --- | --- | --- |
| Mean | 72 ± 13 | 76 ± 16 | 0.03 | 72 ± 9 | 78 ± 9 | < 0.001 |
| Standard Deviation | 6.9 ± 5.0 | 5.7 ± 2.4 | 0.80 | 5.3 ± 2.4 | 6.0 ± 2.0 | 0.17 |

Table 5.2: Global peak strain results for NF  and FG  scans.

|  | No Feedback | Feedback Game | p-value |
| --- | --- | --- | --- |
| Circumferential Strain (%) Radial Strain (%) | -17 ± 2 44 ± 11 | -16 ± 2 40 ± 10 | < 0.001 0.005 |
| Longitudinal Strain (%) | -13 ± 2 | -13 ± 2 | 0.38 |

differences in longitudinal strain between NF and FG acquisitions (13 ± 2% vs 13 ±

2%, p =0.38).

5.5  Discussion

Feedback  of the diaphragm position during CMR has been shown to improve navigator efficiency in adults [68, 69]. This study explored how the use of a feedback game (FG) affects navigator efficiency compared to traditional  no- feedback (NF) acquisitions in children.  The results of the study showed that,  compared to NF, using the FG resulted in 1) substantially improved navigator efficiency (from 33 to

58%); 2) slightly improved SNR; 3) slightly higher mean heart rate; and 4) slightly lower global strain magnitudes. Importantly, these results were not affected by the use of an off-scanner training protocol, which suggests that lengthy, robust training (11 min in our protocol) does not need to be a part of the clinical/imaging workflow for this interface.

5.5.1 Navigator Efficiency

Navigator efficiency  was improved from 33 to 58% by using a FG in children (Figure 5.3a).  This increase in navigator efficiency led to a 43% reduction in the number of heartbeats required to complete a scan. Studies have shown that feedback of the diaphragm position during CMR results in a more reproducible breath-hold position [68, 69, 77], which can lead to improved navigator efficiency. Previous CMR studies have reported that  NF navigator efficiencies can vary from 20 to 45% in adults [71, 68, 69, 74], and we found a comparable NF navigator efficiency of 33% in

children using a conservative dual-navigator acceptance strategy. Visual feedback of the diaphragm  position has been shown to improve end-expiratory navigator efficiency from 45 to 56% [69] and from 42 to 71% with the addition of supplemental oxygen [68] leading to a 20% and 41% reduction in the number of required heartbeats, respectively. With the use of the FG, we found a slightly better improvement of navigator efficiency from 33 to 58% in children without the use of supplemental oxygen. Average navigator efficiency  was not correlated with age (Figure 5.3b). Therefore, children ages eight and older should be able to utilize the FG to effectively improve navigator efficiency compared to conventional NF acquisitions.

Extensive  off-scanner  training  using an MRI-simulator  was  not  necessary  to achieve the observed improvement in navigator efficiency using the FG. Instead, the subjects with minimal training immediately prior to data acquisition had equivalent efficiency  as their extensively-trained counterparts. While this finding might suggest that the chest wall excursion-based training method was ineffective,  it is more likely that the intuitive interface design  was easy to learn and therefore the children did not require much training.   Importantly,  the two  30-beat practice scans provided some degree of training in both cases, which is intuitively  necessary.  Future efforts can optimize that practice time to provide the needed minimal training in the most efficient manner.

5.5.2 SNR

We  found that  using the FG  slightly  improved the SNR of the end-systolic magnitude images of our spiral DENSE sequence  by 5% compared to NF for all images combined (p  = 0.01, Figure  5.5).   This finding contrasts with  previous studies, which reported image quality score using 2 expert reviewers and found that the  use of  diaphragmatic feedback maintained image quality  compared to  NF acquisitions [68, 69]. The difference in image quality is likely sequence dependent. The   previous  studies    were  performed  using   steady-state  free  precession.

Additionally, it is likely that quantitative measurement of SNR is more sensitive at detecting differences  in  image  quality  compared to  subjective  image scoring by expert reviewers.

5.5.3 Heart Rate

A potential negative  finding of this study was the slight  increase in heart rate observed with  the use of the feedback game. To be clear, this difference did not represent  an  increase in  heart  rate  variabilityas  evidenced by  the  comparable standard deviation valuesbut simply a higher baseline value. Such findings are not unprecedented,   as a previous CMR study found a mean heart rate increase of 5 beats/min with use of navigator feedback in adults (compared to our 3 beats/min), and similarly  no differences  in  heart rate variability  between NF  and navigator feedback [68]. A likely reason for this difference is the longer breath-holds performed during the FG, which could have  increased  the heart rate, compared  to relaxed breathing during NF. Another mechanism could be related to stimulation and adrenaline associated with playing the game, compared to the relaxed, passive state associated with NF.

The importance and implications of this potential heart rate difference likely depends  on the imaging application.   While it may mean very little  for purely anatomic evaluations, functional measures,  such as  strains, may be  affected by changing loading conditions and contractility  [78].  To counteract such effects, if undesirable, patients could be coached to relax when playing the game and to not be  too  competitive.  The design of the game could be  modified to enforce such behavior; for example, by programmatically requiring the participant to inhale/exhale after a fixed period of time, or instructing him/her  to periodically take a series of relaxed breaths between cycles of breath-holding.

5.5.4 Strains

We observed small, but statistically significant decreases in global circumferential and radial strains with use of the FG, compared with NF. There was, however, no difference in longitudinal strain.   While these  findings warrant  further study and consideration, the clinical relevance of such small differences (1% for circumferential strain, 4% for radial strain) is likely minimal as they are smaller than previously observed inter-test (±2.0% for circumferential, ±13% for radial) and inter-observer (±1.4% for circumferential, ±14% for radial) 95% limits of agreement  for DENSE [46, 6].

5.5.5 Clinical Implications

Importantly,  the equipment  needed to utilize the FG is minimal and does not directly integrate into an imaging  sequence; it connects externally to the scanner user interface. Due to the minimal equipment needed and non-invasive connection to the MRI scanner, we anticipate that the FG system can be easily adopted at research and clinical sites that perform CMR navigator gating, especially in children. Since navigator efficiency can be increased from 33 to 58%, leading to reduced acquisition times, use of the FG can help improve the clinical feasibility of advanced imaging techniques. While reducing the acquisition time would likely be the most common use of increased navigator efficiency from the FG, the saved time could be allocated to improve image spatial or temporal resolution [68]. Importantly, pre-scan training was not necessary for navigator efficiency improvement with our system,  so clinical and research sites would not have to invest in an MRI simulator environment or spend significant  time training children.  Navigator feedback  has been shown to reduce acquisition time in adults [68], thus, the use of the FG will likely work well in adults also.

Since we only acquired DENSE images for this study, the specific findings are

only definitively relevant for DENSE. However, it is reasonable to expect that these findings are generalizable to many other CMR acquisitions that utilize a respiratory navigator. Possible exceptions include higher resolution applications, such as coronary MR angiography, which may be more sensitive to registration issues.  Further study is needed to test this technique for these applications.

5.5.6 Comparison with Previous Work

A  previous study presented  a respiratory biofeedback  game and continuously adaptive  windowing strategy (CLAWS) to increase navigator efficiency of imaging the thoracic aorta. The authors reported an increase in efficiency in that study from

45 to 56% in adults [69], which represents a smaller magnitude of improvement (25% vs 11%) but a similar end result (58% vs 56%) compared to our study. Although the two  studies are similar, there are distinct differences in design. Most notably, the previous study was in adults; whereas we exclusively  focused on children, based on their limited ability to breath-hold and thus potentially greater need for respiratory navigated  sequences. Additionally, the previous study modified their pulse sequence to allow acquisition of multiple navigator echoes, likely providing a smoother game experience.  We did not modify our cine pulse sequence in our evaluationwe had a single navigator echo  per cardiac cyclein order to  ensure more general clinical applicability.     Collectively,  these  studies demonstrate the  potential  utility   of user-friendly interfaces for improving efficiency and image quality of cardiovascular imaging  sequences using a respiratory navigator in a broad array of patients.

5.5.7 Limitations

This study used a dual-navigator strategy when performing image acquisition. Dual-navigator  strategies have   stricter   data  acceptance criteria  compared to previously used single-navigator strategies [45],  and,  given the  same imaging parameters, will  likely result in lower navigator efficiencies. However, a previous

study using a single-navigator strategy with  navigator feedback  reported similar navigator efficiency results compared to our study.  Therefore, the  use of the  FG with a single- navigator strategy will likely have similar results to this study except that   both  NF  and  FG  acquisitions   may  have  improved  navigator  efficiency compared to a dual-navigator strategy.

The respiratory navigator gating sequence used in this study only measured the diaphragm  position once per cardiac cycle. This low refresh rate can make fine control of the diaphragm position more challenging, especially for participants who may have lower heart rates. Increasing the number of navigator echoes per cardiac cycle could therefore improve performance, but such modifications may not be possible  for all sequences, as is the case for DENSE. Furthermore, even with this limitation, we still found substantial improvement in navigator efficiency when using the FG compared to NF acquisitions.

Due to the randomization of the participants into  the trained and untrained groups, there was no attempt to balance age between groups. Therefore, the average trained participant was about 2 years older than the average untrained participant. We  found that there was no difference in FG navigator efficiency between trained and untrained participants.   Even though there was  an age difference between trained and untrained groups, there was no correlation between  age and navigator efficiency with  the FG (Figure 5.3b); thus, the results of the study apply to all children aged eight to eighteen.

In order to accurately assess the NF navigator efficiency  as it would be in the clinical setting, we did not want  to influence the children’s natural breathing pat- tern. In particular, we did not want the breathing pattern performed during the FG acquisitions to influence the NF breathing pattern. Therefore, NF acquisitions were always performed before FG acquisitions.  Since the order of NF and FG acquisitions was not randomized,  this may have affected the results as  participants may have

become more comfortable  as they spent more time in the MRI scanner. However, performing this randomization likely would have resulted in similar conclusions and we feel that it was important to accurately measure the navigator efficiency of the NF acquisitions.

Due to the potential for patient  movement  or erratic breathing patterns, we utilized a stoppage criterion to attempt to maintain a 30 min protocol length. We observed  eleven  cases which satisfied stoppage criterion and four cases  of patient movement  (one which also satisfied stoppage criterion).   In these participants, we estimated navigator efficiency, SNR, and heart rate from fewer acquisitions than the remaining participants. However,  since we used all of the data that we did acquire for each participant, the computed values are appropriate.

The two 30-heartbeat practice scans were not included in the computation and analysis of navigator efficiency for the FG technique. Their inclusion would only slightly decrease the reported gains in efficiency (for example, if we used the FG to acquire 300 heart beats of actual data, the reduction in scan time would change minimally from 43 to 37% after accounting for the two practice scans); however, it must be noted that the selection of those practice parameters was arbitrary and not optimized. In reality, less training is likely required to familiarize the subject with the interface,  so factoring this specific training design into the analysis is not critical.

We  performed this  study in children with  no significant  past medical history. While we did attempt to recruit from a broad clinical population using recruitment services  at our Center for Clinical and Translational Science,  the population we ultimately studied may not be entirely representative of a standard pediatric clinical population that would routinely undergo cardiac MRI. For example, approximately

25% of  patients  with  tetralogy  of  Fallot  may  have  learning and  behavioral difficulties [79], which may impair their ability to benefit from the feedback game. It is therefore reasonable to expect that  the true benefit of the feedback game in a

standard clinical population will be smaller than what was measured in the current study, but still better than what can be expected without the use of feedback. Even if only half of the patients benefit to the extent  shown in the current  study, the overall navigator efficiency for the clinical population as a whole would still increase from 33% efficiency to 46% efficiency (a 38% relative benefit). Future research will seek to evaluate this in further detail as we  implement  the feedback  game during routine clinical workflows.

5.5.8 Conclusion

Use of a respiratory navigator feedback game designed to engage children during navigator-gated CMR improved navigator efficiency in children from 33 to 58%. This improved efficiency reduces the number of heartbeats and corresponding scan durations by 43%, and is also associated with a 5% increase in SNR for spiral cine DENSE. Pre-scan training on how to use the feedback  game is not necessary  to achieve the improvement in navigator efficiency.  These findings should generalize to all CMR acquisition sequences that utilize a respiratory navigator.

Appendices

APPENDIX A

DIFFERENCES IN SEGMENTAL STRAIN BETWEEN DIFFERENT ACCEPTANCE WINDOW  POSITIONS

Table A.1:  Segmental circumferential strain (%, mean ± standard deviation) from the three acceptance window

positions (minimum, middle, and maximum)  for all subjects combined.

Anterior 	Anteroseptal	Inferoseptal	Inferior	Inferolateral	Anterolateral

|  | Strain | P | Strain | P | Strain | P | Strain | P | Strain | P | Strain | P |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  | Basal |  |  |  |  |  |  |
| Max | -16±5 |  | -14±5 |  | -14±5 |  | -15±5 |  | -19±6 |  | -19±5 |  |
| Mid | -16±3 | 0.99 | -13±5 | 1.0 | -15±5 | 0.95 | -15±5 | 0.91 | -18±5 | 0.88 | -19±5 | 0.76 |
| Min | -15±4 |  | -13±5 |  | -15±4 |  | -15±5 |  | -19±6 |  | -18±4 |  |

Mid-Ventricular

Max 	-17±5

-14±5

-13±5

-17±5

-22±6

-20±6

| Mid | -17±5 | 0.93 | -14±4 | 0.83 | -14±4 | 0.87 | -18±5 | 0.93 | -21±4 | 1.0 | -21±6 | 0.81 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Min | -17±6 |  | -14±4 |  | -13±4 |  | -17±3 |  | -20±5 |  | -21±6 |  |
|  |  |  |  |  |  | Apical |  |  |  |  |  |  |

Max 	-18±5

Mid 	-18±5

0.66

-15±6

-15±5

0.99

-17±6

-17±6

0.79

-20±6

-20±6

0.98

-23±7

-23±6

0.84

-22±6

-22±7

0.99

88

Min 	-18±6	-15±5	-16±6	-21±6	-24±6	-21±7

Table A.2:  Segmental radial strain (%, mean ± standard deviation) from the three acceptance window positions

(minimum, middle, and maximum)  for all subjects combined.

Anterior 	Anteroseptal	Inferoseptal	Inferior	Inferolateral	Anterolateral

|  | Strain | P | Strain | P | Strain | P | Strain | P | Strain | P | Strain | P |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  | Basal |  |  |  |  |  |  |
| Max | 37±20 |  | 36±16 |  | 40±20 |  | 43±23 |  | 47±29 |  | 41±21 |  |
| Mid | 40±21 | 0.78 | 41±18 | 0.61 | 37±14 | 0.53 | 36±21 | 0.77 | 47±31 | 0.94 | 44±24 | 0.64 |
| Min | 38±18 |  | 38±17 |  | 37±18 |  | 40±22 |  | 51±29 |  | 46±28 |  |

Mid-Ventricular

Max 	28±17

33±24

34±15

32±15

36±32

32±17

| Mid Min | 31±19 30±19 | 0.94 | 35±19 36±14 | 0.44 | 31±13 36±16 | 0.88 33±24 31±21 | 0.96 | 38±29 35±24 | 0.83 | 33±18 28±17 | 0.49 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  | Apical |  |  |  |  |  |
| Max | 28±24 |  | 36±37 |  | 41±21 | 41±37 |  | 31±23 |  | 27±16 |  |
| Mid | 23±12 | 0.72 | 31±13 | 0.87 | 39±20 | 0.83 40±20 | 0.20 | 27±17 | 0.84 | 26±15 | 0.69 |
| Min | 25±15 |  | 35±24 |  | 40±27 | 34±24 |  | 29±20 |  | 33±16 |  |

89

APPENDIX B

SEGMENTAL STRAINS FOR  NAVIGATOR GATING

B.1  Peak Segmental Strains for Navigator  Gating and Breath-holds

The myocardium was  divided into  6 segments  for  mid-ventricular short-axis images and 7 segments for four-chamber long-axis  images based on the American Heart Association 17-segment  model.   The peak strain was  computed for each segment and reported using mean and standard deviation of all segments in Table S1.  For four-chamber  images, pixels from the most basal and apical segments (3 segments in total)  were excluded  from analysis in order to remove  the increased noise typically observed in those regions, which also matched the analysis performed for global longitudinal strain.

B.2  Segmental Strain Agreement Between Navigator  Gating and Breath- holds

Peak segmental strains were compared between each  navigator  gating configuration  (retrospective,  prospective,  and  dual)   and  breath-holds  using Bland-Altman  analyses and  coefficient   of  variation  (CoV).  Comparisons were performed using a paired Students t-test.

Table B.1:  Segmental strain results for navigator gating and breath-holds in adults.

Circumferential  Strain (%)

Mean ± Std.  Dev.

Breath-hold	-18 ± 5

Retrospective	-18 ± 5

Prospective	-18 ± 5

Dual	-18 ± 4

Radial Strain (%)

Breath-hold	35 ± 16

Retrospective	34 ± 16

Prospective	42 ± 17

Dual	34 ± 16

Longitudinal  Strain (%)

Breath-hold	-13 ± 4

Retrospective	-14 ± 3

Prospective	-13 ± 3

Dual	-14 ± 4

Table  B.2: 	Segmental strain  agreement  between  navigator  gating  and breath-holds from spiral cine DENSE.

Radial Strain (%)

Longitudinal  Strain (%)

|  | Bias | 95% LoA | CoV (%) | p-value |
| --- | --- | --- | --- | --- |
| Circumferential Strain (%) |  |  |  |  |
| Retrospective–Breath-hold | 0 | ± 5 | 8 | 0.94 |
| Prospective–Breath-hold | 0 | ± 8 | 13 | 0.78 |
| Dual–Breath-hold | -1 | ± 5 | 8 | 0.11 |
| Retrospective–Breath-hold | 0 | ± 25 | 19 | 0.79 |
| Prospective–Breath-hold | 8 | ± 36 | 28 | 0.002* |
| Dual–Breath-hold | -1 | ± 29 | 23 | 0.61 |
| Retrospective–Breath-hold | -1 | ± 7 | 13 | 0.17 |
| Prospective–Breath-hold | 0 | ± 9 | 17 | 0.75 |
| Dual–Breath-hold | -1 | ± 8 | 13 | 0.19 |

* indicates p < 0.05

APPENDIX C

OFF-SCANNER TRAINING PROTOCOL FOR  FEEDBACK GAME C.1  Training Protocol

The  goal-based  training  protocol was   as  follows:	First,  the  children were

instructed to  perform 3 sequential end-expiratory breath-holds to  determine the optimal location for the acceptance window.  Then the children were instructed to complete 9 levels of the FG, which progressed in difficulty.  Difficulty was increased by either 1) decreasing the acceptance window size or 2) increasing the time delay between chest excursion recording  and fish location update. Because the navigator gating sequence only measures the diaphragm position during each heartbeat, children  with   slower heartbeats  may  experience ”delays”  between diaphragm movement  and fish location update.  In order to complete  each level, the children had to acquire 100 points. If all bubbles were acquired in a row with no breaks, each level could be completed in ∼33 seconds.

C.2  Survey Responses

In order to formally measure the enjoyment and response of the children playing the Feedback Game, we asked the children to fill out a post-scan survey that consisted of 7 questions. Those questions and responses are listed below.  In general, most participants 1) found Bubble Gulp to be easy to play; 2) enjoyed playing Bubble Gulp;

3) thought  they were generally getting better as they played; 4) thought  training was/would have been somewhat  helpful; 5) had no comments on how to improve

’Bubble Gulp’; 6) enjoy playing videogames; and 7) play videogames daily.

Question 1 Response

12

10

Number of Children

8

6

4

2

0

Figure C.1:  Question 1. How easy was playing Bubble Gulp?

1: Really easy

2: Easy

3: Neither easy nor difficult

4: Difficult

5: Really Difficult

Question 2 Response

Number of Children

15

10

5

0

Figure C.2:  Question 2. How much did you enjoy playing Bubble Gulp?

1: Really enjoyed it

2: Enjoyed it

3: Neither

4: Did not enjoy it

5: Really did not enjoy it

Question 3 Response

15

Number of Children

10

5

0

Figure C.3:  Question 3. Did you think you were getting better, stayed the same, or were getting worse as you were playing Bubble Gulp at the end of the study compared to when you first tried it?

1: Better

2: Stayed the same

3: Worse

Question 4 Response

15

Number of Children

10

5

0

Figure C.4:  Question 4. We have a pretend MRI scanner where you can learn to play Bubble Gulp before getting into the actual MRI scanner. Do you think using this pretend MRI scanner first would have been/was:

1: Very helpful

2: Somewhat helpful

3: Not helpful

4: A total wast of time

5. Do you have any comments on how to improve Bubble Gulp?

• Mostly ”None”

• ”Make the fish pink”

• ”Make the lines further a part on the screen”

• ”Liked the simple concept and how could control with breathing”

• ”Reverse direction of fish movement with breathing”

• ”Make not as glitchy, (make smoother)”

• ”Make lines move to more comfortable spot to breathe in”

– This subject moved before their last scan

Question 6 Response

20

Number of Children

15

10

5

0

Figure C.5:  Question 6. How much do you enjoy playing videogames?

1: Really enjoy

2: Enjoy

3: Neither

4: Do not enjoy

5: Really do not enjoy

Question 7 Response

Number of Children

15

10

5

0

Figure C.6:  Question 7. How often do you play videogames?

1: Daily

2: 2-3 times per week

3: Weekly

4: 1-2 times per month

5: Seldom to never

Bibliography

[1] CDC.	“Underlying  Cause of Death 1999-2013  on CDC WONDER  Online

Database.”, 2015.

[2] Rosamond W, Flegal K, Friday G, Furie K, Go A, Greenlund K, Haase N, Ho M, Howard V, Kissela B, Kittner  S, Lloyd-Jones D, McDermott M, Meigs J, Moy C, Nichol G, O’Donnell CJ, Roger V, Rumsfeld J, Sorlie P, Steinberger J, Thom T, Wasserthiel-Smoller  S, Hong Y, American Heart Association Statistics C, and Stroke Statistics S. “Heart disease and stroke statistics–2007 update: a report from the American Heart Association Statistics Committee and Stroke Statistics Subcommittee.”  Circulation,  2007.  115(5):e69–171. doi:10.1161/ CIRCULATIONAHA.106.179918.

[3] Stanton T, Leano R, and Marwick TH.  “Prediction of all-cause mortality from global longitudinal speckle strain:  comparison with ejection fraction and wall motion scoring.” Circulation Cardiovascular imaging, 2009. 2(5):356–64.  ISSN

1942-0080. doi:10.1161/CIRCIMAGING.109.862334.

[4] Aletras AH, Ding S, Balaban  RS, and Wen H. “DENSE: displacement encoding with  stimulated echoes  in  cardiac functional MRI.”    J  Magn Reson, 1999.

137(1):247–252. doi:10.1006/jmre.1998.1676.

[5] Haggerty CM, Kramer SP, Binkley CM, Powell DK, Mattingly  AC, Charnigo R, Epstein FH, and Fornwalt BK.    “Reproducibility of cine displacement encoding with stimulated echoes  (DENSE) cardiovascular magnetic resonance for measuring left ventricular strains, torsion, and synchrony in mice.” Journal of  cardiovascular magnetic resonance  :   official  journal  of  the Society for Cardiovascular Magnetic Resonance,  2013. 15(1):71.  ISSN 1532-429X. doi:

10.1186/1532-429X-15-71.

[6] Wehner GJ, Grabau JD, Suever JD, Haggerty CM, Jing L, Powell DK, Hamlet SM, Vandsburger MH, Zhong X, and Fornwalt BK.   “2D  cine DENSE with low encoding frequencies accurately quantifies cardiac mechanics with improved image characteristics.” Journal of cardiovascular magnetic resonance :  official journal of the Society for Cardiovascular Magnetic Resonance, 2015. 17(1):93. ISSN 1532-429X. doi:10.1186/s12968-015-0196-z.

[7] Amundsen BH, Helle-Valle T,  Edvardsen T, Torp H, Crosby J, Lyseggen E, Støylen A, Ihlen H, aC Lima J, Smiseth Oa, and Slørdahl Sa.  “Noninvasive Myocardial Strain Measurement by Speckle Tracking Echocardiography.” Journal of the American College  of Cardiology, 2006.  47(4):789–793. ISSN

07351097. doi:10.1016/j.jacc.2005.10.040.

[8] Hor KN, Gottliebson WM, Carson C, Wash E, Cnota J, Fleck R, Wansapura J,  Klimeczek P,  Al-Khalidi   HR,  Chung ES, Benson DW,  and Mazur  W. “Comparison of Magnetic Resonance Feature Tracking for Strain Calculation With Harmonic Phase Imaging Analysis.” JACC: Cardiovascular Imaging, 2010.

3(2):144–151. ISSN 1936878X. doi:10.1016/j.jcmg.2009.11.006.

[9] Axel L and Dougherty L. 	“MR  imaging of motion with  spatial modulation of magnetization.”	Radiology, 1989.	171(3):841–845.	ISSN 0033-8419.	doi:

10.1148/radiology.171.3.2717762.

[10] Zerhouni EA, Parish DM, Rogers WJ, Yang A, and Shapiro EP. “Human heart: tagging with MR imaging–a method for noninvasive assessment of myocardial motion.” Radiology, 1988. 169(1):59–63.  ISSN 0033-8419. doi:10.1148/radiology.

169.1.3420283.

[11] Pelc LR, Sayre J, Yun K, Castro LJ, Herfkens RJ, Miller  DC, and Pelc NJ. “Evaluation of myocardial motion tracking with cine-phase contrast magnetic resonance  imaging.”   Investigative radiology, 1994.  29(12):1038–1042. ISSN

0020-9996. doi:10.1097/00004424-199412000-00005.

[12] Osman NF, Sampath  S, Atalar E, and Prince JL. “Imaging longitudinal cardiac strain on short-axis images using strain-encoded MRI.”  Magnetic Resonance in Medicine, 2001. 46(2):324–334.  ISSN 07403194. doi:10.1002/mrm.1195.

[13] Aletras AH, Balaban RS, and Wen H.  “High-resolution strain analysis of the human heart with fast-DENSE.”  Journal of magnetic resonance  (San Diego, Calif : 1997), 1999. 140(1):41–57.  ISSN 1090-7807. doi:10.1006/jmre.1999.1821.

[14] Kuijer  JPa, Marcus JT,  G¨otte  MJW,  van Rossum  AC,  and Heethaar RM. “Three-dimensional myocardial strains at end-systole and during diastole in the left ventricle of normal humans.” Journal of cardiovascular magnetic resonance

:  official journal of the Society for Cardiovascular Magnetic Resonance, 2002.

4(3):341–351.  ISSN 1097-6647. doi:Doi10.1081/Jcmr-120013299.

[15] Moore C, Lugo-Olivieri C, McVeigh E, and Zerhouni E.  “Three-dimensional systolic strain patterns in the normal human left ventricle: characterization with tagged MR imaging.” Radiology, 2000. 214(2):453–466.  ISSN 0033-8419.  doi:

10.1148/radiology.214.2.r00fe17453.

[16] Young aa, Kramer CM, Ferrari Va, Axel L, and Reichek N. “Three-dimensional left ventricular deformation in hypertrophic cardiomyopathy.” Circulation, 1994.

90(2):854–867.  ISSN 0009-7322. doi:10.1161/01.CIR.90.2.854.

[17] Feng L, Donnino R, Babb J, Axel L, and Kim  D.   “Numerical and in vivo validation of fast cine displacement-encoded with stimulated echoes (DENSE) MRI  for quantification of regional cardiac function.”   Magnetic resonance  in medicine, 2009. 62(3):682–90.  ISSN 1522-2594. doi:10.1002/mrm.22045.

[18] Nasiraei  Moghaddam A, Saber NR, Wen H, Finn JP, Ennis DB, and Gharib M. 	“Analytical  method to measure three-dimensional  strain patterns in the left ventricle from single slice displacement  data.”   Journal of cardiovascular magnetic resonance : official journal of the Society for Cardiovascular Magnetic Resonance, 2010. 12:33. ISSN 1532-429X. doi:10.1186/1532-429X-12-33.

[19] Donekal S, Ambale-Venkatesh B, Berkowitz S, Wu CO, Choi EY, Fernandes V, Yan R, Harouni Aa, Bluemke  Da, and Lima Ja.  “Inter-study  reproducibility of  cardiovascular magnetic resonance  tagging.”     Journal  of  cardiovascular magnetic resonance : official journal of the Society for Cardiovascular Magnetic Resonance, 2013. 15(1):37. ISSN 1532-429X. doi:10.1186/1532-429X-15-37.

[20] Suever J, Wehner G, Jing L, Powell D, Hamlet  S, Grabau J, Mojsejenko D, Andres K, Haggerty C, and Fornwalt  B.  “Right  Ventricular Strain, Torsion, and Dyssynchrony in Healthy Subjects using 3D Spiral Cine DENSE Magnetic Resonance Imaging.” IEEE Transactions on Medical Imaging, 2017. 1–1. ISSN

0278-0062. doi:10.1109/TMI.2016.2646321.

[21] Liu YL, Riederer SJ, Rossman PJ, Grimm RC, Debbins JP, and Ehman RL. “A Monitoring, Feedback, and Triggering System for Reproducible Breath-Hold MR Imaging.” Magnetic Resonance in Medicine, 1993. 30:507–511.

[22] Wang Y, Christy PS, Korosec FR, Alley MT, Grist TM, Polzin JA, and Mistretta CA.  “Coronary MRI with a Respiratory Feedback Monitor:  The 2D Imaging Case.” Magnetic Resonance in Medicine, 1995. 33:116–121.

[23] Taylor aM, Jhooti P, Wiesmann F, Keegan J, Firmin  DN, and Pennell DJ. “MR  navigator-echo  monitoring of temporal changes  in diaphragm position: implications for MR coronary angiography.”  Journal of magnetic resonance imaging : JMRI, 1997. 7(4):629–36.  ISSN 1053-1807.

[24] Holland AE, Goldfarb JW, and Edelman RR.   “Diaphragmatic and Cardiac Motion during Suspended Breathing: Preliminary Experience and Implications for Breath-hold MR Imaging.” Radiology, 1998. 209(2):483–489.

[25] Fischer RW, Botnar RM, Nehrke K, Boesiger P, Manning WJ, and Peters DC. “Analysis of residual coronary artery motion for breath hold and navigator approaches  using real-time coronary MRI.”    Magnetic resonance  in medicine

:  official journal of the Society of Magnetic Resonance in Medicine / Society of Magnetic Resonance in Medicine, 2006. 55(3):612–8. ISSN 0740-3194.  doi:

10.1002/mrm.20809.

[26] Slomka PJ, Fieno D, Ramesh A, Goyal V, Nishina H, Thompson LE, Saouaf R, Berman DS, and Germano G.  “Patient motion correction for multiplanar, multi-breath-hold cardiac cine MR imaging.”  J Magn Reson Imaging, 2007.

25(5):965–973. doi:10.1002/jmri.20909.

[27] Swingen C, Seethamraju RT, and Jerosch-Herold M. “An approach to the three- dimensional display of left ventricular function and viability  using MRI.”  Int J Cardiovasc Imaging, 2003. 19(4):325–336.

[28] Wang Y,  Riederer SJ, and Ehman RL.   “Respiratory Motion of the Heart: Kinematics and the Implications for the Spatial Resolution in Coronary Imaging.” Magnetic Resonance in Medicine, 1995. 33:713–719.

[29] McLeish K, Hill  DLG, Atkinson D, Blackall JM, and Razavi R.  “A  study of the motion and deformation of the heart due to respiration.” Medical Imaging, IEEE Transactions on, 2002. 21(9):1142–1150.  doi:10.1109/tmi.2002.804427.

[30] Hamlet SM, Haggerty CM, Suever JD, Wehner GJ, Grabau JD, Andres KN, Vandsburger MH, Powell DK, Sorrell VL, and Fornwalt BK.  “An  interactive videogame designed to  improve  respiratory navigator efficiency in  children undergoing cardiac magnetic resonance.” Journal of Cardiovascular Magnetic Resonance, 2016. 18(1):54. ISSN 1532-429X. doi:10.1186/s12968-016-0272-z.

[31] Zhong X, Spottiswoode BS, Cowart Ea, Gilson WD, and Epstein FH. “Selective suppression of artifact-generating echoes  in cine DENSE using through-plane dephasing.” Magnetic resonance in medicine : official journal of the Society of Magnetic Resonance in Medicine / Society of Magnetic Resonance in Medicine,

2006. 56(5):1126–31.  ISSN 0740-3194. doi:10.1002/mrm.21058.

[32] Kim D, Gilson WD, Kramer CM, and Epstein FH. “Myocardial tissue tracking with  two-dimensional  cine displacement-encoded  MR  imaging:  development and initial  evaluation.”  Radiology, 2004. 230(3):862–871. doi:10.1148/radiol.

2303021213.

[33] Hamlet  SM,  Haggerty  CM,  Suever JD,  Wehner GJ,  Andres KN,  Powell DK,  Zhong X, and Fornwalt BK.    “Optimal  configuration of respiratory navigator gating for the quantification of left  ventricular strain using spiral cine displacement  encoding with stimulated echoes  (DENSE) MRI.”   Journal of magnetic resonance  imaging :  JMRI, 2016. ISSN 1522-2586.   doi:10.1002/ jmri.25389.

[34] Gilliam AD and Suever JD. “DENSEanalysis.”, 2016.

[35] Suever  JD,  Wehner GJ,  Haggerty CM,  Jing  L,  Hamlet SM, Binkley  CM, Kramer SP, Mattingly AC, Powell DK, Bilchick KC, Epstein FH, and Fornwalt BK. 	“Simplified  post processing  of  cine DENSE cardiovascular magnetic resonance for quantification of cardiac mechanics.” Journal of cardiovascular magnetic resonance : official journal of the Society for Cardiovascular Magnetic Resonance, 2014. 16:94. ISSN 1532-429X. doi:10.1186/s12968-014-0094-9.

[36] Spottiswoode BS, Zhong X, Hess aT, Kramer CM, Meintjes EM, Mayosi BM, and Epstein FH. “Tracking myocardial motion from cine DENSE images using spatiotemporal  phase unwrapping and temporal fitting.”  IEEE transactions on medical imaging, 2007. 26(1):15–30. ISSN 0278-0062.  doi:10.1109/TMI.2006.

884215.

[37] Dytham C. Choosing and Using Statistics: A Biologist’s Guide, 3rd Edition. 3rd edition, 2011.

[38] Bland J and Altman D.  “Statistical methods for assessing agreement between two methods of clinical measurement.”  Lancet, 1986. 1:307–310.

[39] Hamlet SM, Haggerty CM, Suever JD, Wehner GJ, Andres KN, Powell DK, Charnigo RJ, and Fornwalt BK.   “Using a respiratory navigator significantly reduces variability when quantifying left ventricular torsion with cardiovascular magnetic resonance.” Journal of Cardiovascular Magnetic Resonance,  2017.

19(1):25. ISSN 1532-429X. doi:10.1186/s12968-017-0338-6.

[40] Pahlm US, Ubachs JFA, Heiberg E, Engblom H, Erlinge D, G¨otberg  M, and Arheden H. “Regional wall function before and after acute myocardial infarction; an experimental study in pigs.” BMC Cardiovascular  Disorders, 2014. 14(1):1–

10. ISSN 1471-2261.

[41] Ru¨ssel IK, G¨otte MJW, Bronzwaer JG, Knaapen P, Walter J, and Rossum ACV. “Left Ventricular Torsion.” JACC: Cardiovascular  Imaging, 2011. 2(5):648–655. ISSN 1936-878X. doi:10.1016/j.jcmg.2009.03.001.

[42] G¨otte MJW, Germans T, Ru¨ssel IK, Zwanenburg JJM, Marcus JT, van Rossum AC, and van Veldhuisen DJ.   “Myocardial Strain and Torsion Quantified by Cardiovascular Magnetic Resonance  Tissue Tagging. Studies in  Normal and Impaired Left Ventricular Function.”    Journal of the American College of Cardiology, 2006. 48(10):2002–2011.  ISSN 07351097.  doi:10.1016/j.jacc.2006.

07.048.

[43] Kowallick JT,  Morton  G, Lamata P, Jogiya R, Kutty S, Lotz J, Hasenfuß G, Nagel E, Chiribiri  A, and Schuster A.  “Inter-study  reproducibility of left ventricular torsion and torsion rate quantification using MR myocardial feature tracking.”  Journal of magnetic resonance imaging :  JMRI, 2016. 43:128–137. ISSN 1522-2586. doi:10.1002/jmri.24979.

[44] Kaku K, Takeuchi M, Tsang W, Takigiku K, Yasukochi  S, Patel AR, Mor-Avi V, Lang RM, and Otsuji Y. “Age-related normal range of left ventricular strain and torsion using three-dimensional speckle-tracking echocardiography.”  Journal of the American Society of Echocardiography, 2014. 27(1):55–64. ISSN 10976795. doi:10.1016/j.echo.2013.10.002.

[45] Zhong X, Spottiswoode BS, Meyer CH, Kramer CM, and Epstein FH. “Imaging three-dimensional myocardial mechanics using navigator-gated  volumetric spiral cine DENSE MRI.”   Magn Reson Med, 2010. 64(4):1089–1097.   doi:10.1002/ mrm.22503.

[46] Wehner GJ, Suever JD, Haggerty CM, Jing L, Powell DK, Hamlet SM, Grabau JD, Mojsejenko WD, Zhong X, Epstein FH, and Fornwalt BK.   “Validation of in  vivo  2D displacements  from spiral cine DENSE at  3T.”    Journal of Cardiovascular Magnetic Resonance, 2015. 17(1):1–11. ISSN 1532-429X. doi:

10.1186/s12968-015-0119-z.

[47] Phan TT,  Shivu GN, Abozguia K, Gnanadevan M, Ahmed I, and Frenneaux

M.  “Left  ventricular torsion and strain patterns in heart failure with normal

ejection fraction are similar to  age-related changes.”  European  Journal of Echocardiography, 2009.    10(6):793–800.   ISSN 1532-2114.   doi:10.1093/ ejechocard/jep072.

[48] Sorger JM, Wyman BT, Faris OP, Hunter WC, and McVeigh ER. “Torsion of the left ventricle during pacing with MRI tagging.” J Cardiovasc Magn Reson,

2003. 5(4):521–530.

[49] Grothues F,  Smith  GC,  Moon JCC,  Bellenger NG,  Collins P,  Klein  HU, and Pennell DJ.  “Comparison of interstudy reproducibility of cardiovascular magnetic resonance with two-dimensional echocardiography in normal subjects and in  patients with  heart failure or left  ventricular hypertrophy.”    Am  J Cardiol, 2002. 90(1):29–34. ISSN 0002-9149. doi:http://dx.doi.org/10.1016/ S0002-9149(02)02381-0.

[50] Broussard JL, Nelson MD, Kolka CM, Bediako IA, Paszkiewicz RL, Smith L, Szczepaniak EW, Stefanovski D, Szczepaniak LS, and Bergman RN.  “Rapid development of cardiac dysfunction in a canine model of insulin resistance and moderate obesity.”  Diabetologia, 2016. 59(1):197–207.  ISSN 0012-186X. doi:

10.1007/s00125-015-3767-5.

[51] Kramer SP, Powell DK,  Haggerty CM,  Binkley CM,  Mattingly  AC,  Cassis La, Epstein FH, and Fornwalt BK.   “Obesity reduces left ventricular strains, torsion, and synchrony in mouse models: a cine displacement  encoding with stimulated echoes (DENSE) cardiovascular  magnetic resonance study.” Journal of  cardiovascular magnetic resonance  :   official  journal  of  the Society for Cardiovascular Magnetic Resonance,  2013.  15:109.  ISSN 1532-429X.  doi:

10.1186/1532-429X-15-109.

[52] Oxenham HC, Young AA, Cowan BR, Gentles TL, Occleshaw CJ, Fonseca CG, Doughty RN, and Sharpe N.  “Age-Related Changes in Myocardial Relaxation Using Three-Dimensional Tagged Magnetic Resonance  Imaging.”   Journal of Cardiovascular Magnetic Resonance, 2003. 5(3):421–430.  ISSN 10976647.  doi:

10.1081/JCMR-120022258.

[53] Setser RM, Kasper JM, Lieber ML, Starling RC, McCarthy PM, and White RD. “Persistent abnormal left ventricular systolic torsion in dilated cardiomyopathy after partial  left  ventriculectomy.”   Journal of Thoracic and Cardiovascular Surgery, 2003. 126(July):48–55. ISSN 00225223.  doi:10.1016/S0022-5223(03)

00050-3.

[54] Garot J, Pascal O, Di´ebold B, Derumeaux G, Gerber BL, Dubois-Rand´e JL, Lima JaC, and Gu´eret  P.  “Alterations  of systolic left ventricular twist  after acute myocardial infarction.”   American journal of physiology Heart and circulatory physiology, 2002. 282(1):H357–H362. ISSN 0363-6135.   doi:10.1152/ajpheart.

00136.2001.

[55] Kowallick  JT,  Lamata P,  Hussain ST, Kutty   S, Steinmetz M,  Sohns JM, Fasshauer  M, Staab W, Unterberg-Buchwald C, Bigalke  B, Lotz J, Hasenfuß G, and Schuster A. “Quantification of left ventricular torsion and diastolic recoil using cardiovascular  magnetic resonance myocardial  feature tracking.” PloS one,

2014. 9(10):e109164.  ISSN 1932-6203. doi:10.1371/journal.pone.0109164.

[56] Reyhan ML, Wang Z, Kim HJ, Halnon NJ, Finn JP, and Ennis DB. “Effect of free-breathing on left ventricular rotational mechanics in healthy subjects and patients with duchenne muscular dystrophy.”  Magnetic resonance in medicine,

2017. 77(2):864–869.  ISSN 1522-2594. doi:10.1002/mrm.26137.

[57] Ernande L, Thibault  H, Bergerot C, Moulin P, Wen H, Derumeaux G, and Croisille P. “Systolic Myocardial Dysfunction in Patients with Type 2 Diabetes Mellitus :  Identification at MR Imaging with Cine Displacement.”  Radiology,

2012. 265(2):402–409.  doi:10.1148/radiol.12112571/-/DC1.

[58] Aletras AH, Ingkanisorn WP, Mancini C, and Arai AE. “DENSE with SENSE.” Journal of magnetic resonance (San Diego, Calif : 1997), 2005. 176(1):99–106. ISSN 1090-7807. doi:10.1016/j.jmr.2005.05.010.

[59] Young Aa, Li  B, Kirton  RS, and Cowan BR.   “Generalized spatiotemporal myocardial strain  analysis for  DENSE and SPAMM  imaging.”     Magnetic resonance in medicine, 2012. 67(6):1590–9.  ISSN 1522-2594. doi:10.1002/mrm.

23142.

[60] Kar J, Knutsen AK, Cupps BP, Zhong X, and Pasque MK. “Three-dimensional regional strain computation method with displacement encoding with stimulated echoes (DENSE) in non-ischemic, non-valvular dilated cardiomyopathy patients and healthy subjects validated by tagged MRI.”  Journal of magnetic resonance imaging : JMRI, 2014. 00:1–11.  ISSN 1522-2586. doi:10.1002/jmri.24576.

[61] Auger DA,  Zhong X,  Epstein FH,  and Spottiswoode  BS.   “Mapping  right ventricular myocardial mechanics using 3D cine DENSE cardiovascular magnetic resonance.” J Cardiovasc Magn Reson, 2012. 14:4. doi:10.1186/1532-429X-14-4.

[62] Wehner GJ, Suever JD, Haggerty CM, Jing L, Powell D, Zhong X, Epstein FH, and Fornwalt BK. “High resolution cine displacement encoding with stimulated echoes  (DENSE) at 3T with  navigator feedback for quantification of cardiac mechanics.” Journal of Cardiovascular Magnetic Resonance, 2014. 16(Suppl

1):P48. ISSN 1532-429X. doi:10.1186/1532-429X-16-S1-P48.

[63] Stuber  M,  Spiegel  MA,  Fischer SE, Scheidegger  MB,  Danias PG, Pedersen EM, and Boesiger P.  “Single breath-hold slice-following CSPAMM myocardial tagging.” MAGMA, 1999. 9(1-2):85–91.

[64] Swoboda PP, Larghat A, Zaman A, Fairbairn TA, Motwani M, Greenwood JP, and Plein S.  “Reproducibility  of myocardial strain and left ventricular twist measured using complementary spatial modulation of magnetization.” Journal

of magnetic resonance  imaging :  JMRI, 2014. 39(4):887–94. ISSN 15222586. doi:10.1002/jmri.24223.

[65] Moody WE, Taylor RJ, Edwards NC, Chue CD, Umar F, Taylor TJ,  Ferro CJ,  Young AA,  Townend JN,  Leyva F,  and Steeds RP.    “Comparison of magnetic resonance feature tracking for systolic and diastolic strain and strain rate calculation with  spatial modulation of magnetization imaging analysis.” Journal of magnetic resonance imaging : JMRI, 2015. 41(4):1000–1012.  ISSN

1522-2586 (Electronic). doi:10.1002/jmri.24623.

[66] Sigfridsson A, Haraldsson H, Ebbers T, Knutsson H, and Sakuma H.  “In  vivo SNR in DENSE MRI; temporal and regional effects of field strength, receiver coil sensitivity and flip angle strategies.”  Magnetic resonance  imaging, 2011.

29(2):202–8.  ISSN 1873-5894. doi:10.1016/j.mri.2010.08.016.

[67] Gudbjartsson H and Patz S. “The Rician Distribution of Noisy MRI Data.” Magn

Reson Med, 1995. 34(6):910–914. ISSN 15378276. doi:10.1016/j.biotechadv.2011.

08.021.Secreted.

[68] Feuerlein S, Klass O, Pasquarelli A,  Brambs HJ, Wunderlich A,  Duerk JL, Aschoff AJ, and Hoffmann MHK.   “Coronary MR Imaging:  Navigator Echo Biofeedback  Increases Navigator Efficiency-Initial Experience.”  Acad Radiol,

2009. 16(3):374–379.  doi:DOI10.1016/j.acra.2008.08.015.

[69] Jhooti P, Haas T, Kawel N, Bremerich J, Keegan J, and Scheffler  K.   “Use of respiratory biofeedback  and CLAWS for increased navigator efficiency for imaging the thoracic aorta.”  Magn Reson Med, 2011. 66(6):1666–1673.   doi:

10.1002/mrm.22945.

[70] Peters DC, Nezafat  R, Eggers H, Stehning C, and Manning WJ.   “2D  free- breathing dual navigator-gated cardiac function validated against the 2D breath- hold acquisition.” J Magn Reson Imaging, 2008. 28(3):773–777. ISSN 1878-5832. doi:10.1002/jmri.21417.

[71] Abd-Elmoniem KZ, Obele CC, Sibley CT, Matta JR, Pettigrew RI, and Gharib

AM. “Free-breathing single navigator gated cine cardiac magnetic resonance at 3

T: feasibility study in patients.” J Comput Assist Tomogr, 2011. 35(3):382–386. doi:10.1097/RCT.0b013e31821b0ade.

[72] Potthast S, Mitsumori L, Stanescu LA, Richardson ML, Branch K, Dubinsky TJ, and Maki JH.  “Measuring aortic diameter with different MR techniques: Comparison of three-dimensional  (3D)  navigated steady-state  free-precession (SSFP), 3D contrast-enhanced magnetic resonance angiography  (CE-MRA), 2D T2 black blood, and 2D cine SSFP.” Journal of Magnetic Resonance Imaging,

2010. 31(1):177–184.  ISSN 10531807. doi:10.1002/jmri.22016.

[73] Markl M, Kilner PJ, and Ebbers T.  “Comprehensive 4D velocity mapping of the heart and great vessels by cardiovascular magnetic resonance.” J Cardiovasc Magn Reson, 2011. 13:7. doi:10.1186/1532-429X-13-7.

[74] Wang Y, Rossman PJ, Grimm Rc, Riederer SJ, and Ehman RL.  “Navigator- Echo-based Real-Time  Respiratory  Gating  and  Triggering  for  Reduction of  Respiration Effects in  Three-dimensional Coronary  MR  Angiography.” Radiology, 1996. 198:55–60.

[75] Harte JM, Golby CK, Acosta J, Nash EF, Kiraci E, Williams MA, Arvanitis TN, and Naidu B. “Chest wall motion analysis in healthy volunteers and adults with cystic fibrosis using a novel Kinect-based motion tracking system.” Medical

& biological engineering & computing, 2016. ISSN 1741-0444 (Electronic). doi:

10.1007/s11517-015-1433-1.

[76] Heß M, Bu¨ther F, Gigengack F, Dawood M, and Sch¨afers KP.  “A  dual-Kinect approach to determine torso surface motion for respiratory motion correction in PET.”  Medical Physics, 2015. 42(5):2276–2286.  ISSN 0094-2405. doi:10.1118/

1.4917163.

[77] Kim T, Pollock S, Lee D, O’Brien R, and Keall P.  “Audiovisual biofeedback improves   diaphragm motion  reproducibility  in  MRI.”      Med  Phys,  2012.

39(11):6921–6928. doi:10.1118/1.4761866.

[78] Boettler P, Hartmann M, Watzl K, Maroula E, Schulte-moenting  J, Knirsch W, and Dittrich  S. “Heart Rate Effects on Strain and Strain Rate in Healthy Children.”  J Am Soc Echocardiogr,  2005. 18:1121–1130.  ISSN 08947317.  doi:

10.1016/j.echo.2005.08.014.

[79] Piran S, Bassett AS, Grewal J, Swaby JA, Morel C, Oechslin EN, Redington AN, Liu PP, and Silversides CK. “Patterns of cardiac and extracardiac anomalies in adults with tetrology of Fallot.”  American heart journal, 2011. 161(1):131–137. ISSN 1535-7228. doi:10.1016/j.ahj.2010.09.015.